Mover system for BPM
Bino
BPM energy spectrometer meeting
16 november 2006 - UCL

---
Total space available between girder and beamline : 17” ( 431.8 mm )
Space needed :
T0 ( nominal ) :	6.36” ( 161.5 mm ) ( travel range : +/- 0.33” )
BAZ stage ( nominal ) :	4.375“ ( 111.1 mm ) ( travel range : +/- 0.095 “
TM-200 stage :	2.500
plate for BPM support : 1 cm = 0.4 inch
tight space budget : ask SLAC to come up with something different than T0 ?
Asked Ray...
Btw.. wrong orientation for T0 supports in this rough sketch

---
Solution for space budget... how to gain 1 in ?
Screw BAZ directly into the T0 stage leaving out it's top plate... should be possible...
make 3 holes in side of BAZ, 0.5 in from the bottom and 3/8-16 thread 0.62 in deep

---
Horizontal stage : NEAT TM-200 Precision Grade
2” full travel range, 15 um accuracy, 3 um repeatability, 35 kg load
23 frame motor mount
2Mm precision leadscrew
0.1 um linear encoder option ( Reinshaw RGH type )
Need to receive full specs from Heason
Imperial mounting holes on bottom of the stage to match the custom pattern
on top of the BAZ stage ( .209”-32 [~ 5.3mm]) in 3”x3” square on center of stage
Mechanical limit switches ( no default hall type ) : switch off the stepper motor
via controller
Need to receive full specs on type + mounting location of these switches
Screwed ( not glued ) into position at +/- 7 mm from the stage's nominal position
Machine protection, put hard mechanical stops at +/- 9 mm so beam can
Never drill into the copper of BPM, even if beam +5mm on one side and BPM
moved to other side
When in energy measurement position ( +5 mm ), then still +/- 2mm left
For calibration ( sufficient ! )
8 Custom metric M6 holes on the top face of the stage in same pattern as
on next page...
Both stepper motor and limit switches/encoder are 15pin D-type connector

---
.190 “ screws ( ~ M5 )
--> asked for metric M6 holes
These holes don't exist for TM-200
Limit switch + encoder
15 pin D type connectors
.209” screws to fix to BAZ stage
~5.3 mm
So hole pattern in bottom is 3inch by 3 inch square

---
Vertical stage : NEAT BAZ UltraStiff Vertical Elevation stage
5 mm travel, 10 um accuracy, 2 um repeatability, 50 kg load
23 frame motor mount
2Mm pitch precision leadscrew
0.5 um linear encoder option ( Reinshaw RGH type )
Need to receive exact specs from Heason
Imperial mounting holes at bottom to fix to T0 stage ( or eventually something else )
Custom imperial hole pattern on top to match hole pattern on bottom of TM-200
Mechanical end switches on side of the stage, around full 5 mm travel range
Need to receive exact specs from Heason
8 Holes for mounting to the bottom are .209 inch ( ~5.3 mm ), 1 inch of thread,
counter sunk, with .330 inch of screw head width, pattern see next image.
Can be mounted by moving the table -> holes become visible....

---
Use these holes to put mechanical stops

---
+48 V
Stepper motors : 2 Baldor DSM 23F-248-M motors
Can provide enough torque for both TM-200 and BAZ to lift 50 kg
Micro stepping drive up to 51200 steps per revolution,
Coupling matches 23 frame mount
-M : manual knob present
Connecting wires go free, but Heason
provides cable from stages to controller
Need intial setup, cable & software delivered by Heason ( number of steps etc... )
200 W, 48 V PSU quoted, can do 220V as well as 110V

---
Magnetic shielding... consider mu-metal for stepper motors ???
In any case use mechanic end switches and not hall effect ones to be sure...
Simulation of 10D37 magnet : 37 inches long ( 93.98 cm ),
0.4 Gauss at 100 cm from magnet center
Should be fine for steppers... maybe find out where exactly BPM located in z,
but probe more then 0.5 m away...
Half magnet length : 466.9 mm

---
Motion control : Baldor E100 ethernet controller
Need firmware written in MINT
workbench + manual freely
downloadable from the net

---
Able to control 4 stepper axes, has USB and RS232 ports for uploading the firmware and
ethernet for control and monitoring. Power supply quoted can do 220V and 110V.
Write a program in MINT ( based upon BASIC) that checks/fills registers on communications
processor and send then commands over ethernet which will put/read commands in the
registers
Need to have a closer look, but according to the Baldor software engineers, it should be
fairly straightforward what we want to do. Heason proposed that someone goes there for
half a day free of charge after having got more familiar with this MINT language.
Response :
This all sounds very straight forward, the activeX supplied with Mint
Workbench can connect to a controller over Ethernet TCP/IP, and then use
any "ICM" instructions that are required, this can be in the form of
individual MOVE instructions or COMMS commands which just write to
registers and are then interpreted by a Mint program installed on the
Mint controller. Providing you are only using stepper axes and not
Ethernet EPL servo axes then there is no complication in doing this at
all.
I suggest the easiest way is to write a Mint program to receive and load
data into comms registers which are accessed over Ethernet from a
LabView program using the our ActiveX control.
Functions could include Move To Home, Move To Position etc and a back
ground task could copy the encoder counter values into other registers
for feeding back to the LabView program.

---
Heason provides a full set of cables up to 3 m needed to setup the steppers,
connect the motors and the encoders to the controller and a USB cable to
program the controller.
This 3 m is sufficient to put the controller somewhere in ESA incapsulated in lead
and is convenient to test here in the lab.
Delivery time : stages 10-12 weeks, everything else can be delivered
in 4-6 weeks, so we can start familiarizing ourselves with the control -> need to
put it on the order to have separate invoices.
Don't order the fully integrated rack mounted version they propose. Also don't order
straightaway the support ( expensive ). They offer half a day free of charge support at
Heason ( Horsham, south-east of Guildford ). And if that doesn't suffice, they offer
£450/day off site, £650/day on site...
Purchase PC for control of movers, aiming at £350 - £400
Lot of PCI slots ( 4 ), at least 1 RS232 port... for the rest standard config
Found proper motherboard on www.microdirect.co.uk ( but not UCL contractor )
P4, bus speed 1066 Mhz, 1 serial port, 4 free PCI slots,
Graphics card -> AGP and 1 PCI Express port
4 USB-2 ports
Contacted Dell to wether they have something...
Buy mini-tower to allow space for full size PCI cards ( like NI-DAQ cards...)

---
Talk about design of L-bracket to support BPM