## Slide 1: Oscilloscope Data Analysis and LED Tests
Anastasia Basharina-Freshville
(University College London)

## Slide 2
Oscilloscope Data Analysis: 24th -25th November 2016
Use a low rate 30 kHz file to inspect waveforms and train the Χ2 value to cut on
For a 30 kHz rate file taken with the digitiser the ADC μ is:
  6127 – 6141 in ADC mode (run 001)
  6206 – 6231 in waveform mode (run 003)
  for a gate of 150 ns
For 30 kHz run with the digitiser in waveform mode (run 003) the integration starts at 47 ns and finishes at 197 ns
Integration start
Integration end
150 ns integration gate

## Slide 3
Oscilloscope Data Analysis: 24th -25th November 2016
For 30 kHz run taken with the oscilloscope (1.98mmCol_RateTest_30kHz):
  The waveforms start at ~ 200 ns
  Therefore to compare to a ~ 150 ns gate as taken with the digitiser use an integration gate of 0 – 350 ns
  Tests have shown that including more of the baseline (from 0 to 200 ns) in the integration does not shift the ADC mean value (see Slide 5)
Integration start
Integration end
350 ns integration gate, equivalent to 150 ns as used for the digitiser

## Slide 4
Look at the χ2/NDF distribution for 30 kHz run taken with the oscilloscope (1.98mmCol_RateTest_30kHz):
Use χ2/NDF < 25!
Χ2/NDF Cut Value to Use for Scope Data
Keep
Discard

## Slide 5
ADC Equivalent Spectra for Different Integration Gates
1.98 mm Ø collimator
30 kHz rate file
Using 0 – 350 ns gate and 190 – 350 ns gate gives the same result.
Therefore it’s okay to use 0 – 350 ns and include “extra baseline”, in case there is any variation in the waveform start.
Using 0 – 300 ns gate gives a μ closer to that obtained with the digitiser:
  ADC mode: 6127 - 6141
  Waveform mode: 6206 - 6231

## Slide 6
30 kHz Rate ADC Spectra with Digitiser and Scope
Why do we see an improved ΔE/E when using the scope?

## Slide 7
Reminder: ADC Spectra for Different Rates from the Digitiser Data (Nov 2016)
As rate increases, the μ also increases

## Slide 8
ADC Equivalent Specta for Different Rates: 300 ns gate, no Χ2/NDF cut
We see the same trend in the shifting of the μ.
At higher rates, the ΔE/E gets worse. This is the opposite for the digitiser collected waveform data.
Why the strange distribution at 300 kHz?
The μ becomes lower at 400 kHz – but is this the usual effect we see at higher rates?

## Slide 9
ADC Equivalent Specta for Different Rates: 300 ns gate, Χ2/NDF < 25
Using the χ2/NDF cut does not seem to have the same effect as with the digitiser waveform data, where the ~ 200 kHz data μ is reduced to be ~ that of the rest of the data when using the χ2/NDF cut.
The “peak” part of the 300 kHz distribution is not improved by the χ2/NDF cut.

## Slide 10
ADC Equivalent Specta for Different Rates: 350 ns gate, no Χ2/NDF cut
Similar picture as for the 300 ns gate with slightly higher ADC μs.

## Slide 11
ADC Equivalent Specta for Different Rates: 350 ns gate, Χ2/NDF < 25
Similar picture as for the 300 ns gate with slightly higher ADC μs.

## Slide 12
Oscilloscope Data Analysis: 2nd – 3rd August 2016

## Slide 13
Χ2/NDF Cut Value to Use for August Scope Data
Look at the χ2/NDF distribution for 25 kHz run taken with the oscilloscope (1.98mmCol_NoAbs_25kHzRate):
Use χ2/NDF < 25!
Keep
Discard

## Slide 14
Oscilloscope Data Analysis: 2nd -3rd November 2016
For 25 kHz run taken with the oscilloscope (1.98mmCol_NoAbs_25kHzRate):
  The waveforms start at ~ 100 ns
  Therefore to compare to a ~ 150 ns gate as taken with the digitiser use an integration gate of 0 – 250 ns
  Previous tests have shown that including more of the baseline in the integration does not shift the ADC mean value (see Slide 5)
Integration start
Integration end
250 ns integration gate, equivalent to 150 ns as used for the digitiser

## Slide 15
25 kHz Rate ADC Spectra with Digitiser and Scope
Do we again see an improved ΔE/E when using the scope?
  Difficult to judge as we do not reproduce the same μ?
The μ of this data does not seem to be as well matched as November 2016 data.
We want to keep at least a 150 ns gate, so use a gate of 0 – 250 ns for scope data integration.
Including 100 ns of baseline at the start of the integration should not change the μ.

## Slide 16
Reminder: ADC Spectra for Different Rates from the Digitiser Data (Aug 2016)
As rate increases, the μ decreases

## Slide 17
ADC Equivalent Specta for Different Rates: 250 ns gate, no Χ2/NDF cut
We see the same trend in the shifting of the μ.
At higher rates, the ΔE/E gets worse. This is the opposite for the digitiser collected waveform data.
Usual “nonsensical” looking data at 500 kHz

## Slide 18
ADC Equivalent Specta for Different Rates: 250 ns gate, Χ2/NDF < 25
Using the χ2/NDF gets rid of events beyond the peak.
The 500 kHz distribution is not improved by the χ2/NDF cut.

## Slide 19
Conclusions
Both the digitiser and the scope data show the same pattern of the μ depending on the rate
When using the digitiser:
  Energy resolution “improves” with rates
When using the scope:
  Energy resolution “worsens” with rates. This is probably more sensible?
Some “strange” looking data taken with the scope: 300 kHz November 2016 data
From these tests do we trust the scope enough to only take scope data?
  Faster to acquire than with the digitiser.
  I would suggest to continue using both for the time being.
Aside comment:
  Processing the scope data seems to be faster than processing the digitiser data. However, there could have been some updates made to the batch farm over Christmas and the speed improvement could be due to that.
What else can we look at with this data?
  Direct comparison of files taken with the 50 Ω splitter, but need to decide what gate to use. Perhaps from 0 to end of waveform for both for an apple to apple comparison?
Look at the “peak count” scope data in order to confirm rates measured with scope

## Slide 20
LED Tests at UCL

## Slide 21
LED Test Procedure
Testing procedure:
  Use pulser to pulse LED at a mean corresponding to that of a low rate run without any absorbers at Clatterbridge:
    ~ 2.1 V, 20 ns width
  Warm LED up for ~ 2 hours to stabilise before data taking.
  Take data with the digitiser using the “trigger out” as an external trigger input for the digitiser.
  Take data at different “constant” rates, from ~ 100 Hz to 1 MHz.
  Take one run whilst varying rates in “real time”.

## Slide 22
LED Tests: Hamamatsu Base (Comparison to August 2016 Clatterbridge Data)
LED pulsed at 2.11 V, PMT HV: - 900 V
Gaussian distribution as expected
The same trend as we see with Clatterbridge data – as the rate increases, the μ decreases (see Slide 16).
A large change occurs between a rate of 1 kHz and 10 kHz – this happens at a lower rate than we would expect – requires further investigation.

## Slide 23
LED Tests: Hamamatsu Base (Comparison to August 2016 Clatterbridge Data)
Run taken whilst varying rates
Confirmation of what we see at Clatterbridge or could this be caused by an unstable LED?
To untangle the two:
  Deliver pulser signal directly to the cathode of the PMT base

## Slide 24
LED Tests: UCL Soldered Base (Comparison to November 2016 Clatterbridge Data)
LED pulsed at 2.13 V, PMT HV: - 900 V
Gaussian distribution as expected
“Significant” deviation from the mean starts at a rate of ~ 200 kHz.
The same trend as we see with Clatterbridge data – as the rate increases, the μ increases (see Slide 7).
The largest mean is for a rate of 500 kHz.
For rates > 500 kHz the mean begins to drop – the usual high rate effect that we see.

## Slide 25
LED Tests: UCL Soldered Base (Comparison to November 2016 Clatterbridge Data)
LED pulsed at 2.13 V, PMT HV: - 900 V
Gaussian distribution as expected
The mean appears to be more consistent at lower rates.
But why the change to lower means from 1 kHz to 5 -50 kHz?

## Slide 26
LED Tests: UCL Soldered Base (Comparison to November 2016 Clatterbridge Data)
Run taken whilst varying rates
Confirmation of what we see at Clatterbridge or could this be caused by an unstable LED?
To untangle the two:
  Deliver pulser signal directly to the cathode of the PMT base