## Slide 1: Clatterbridge 2nd – 3rd August 2016 Waveform Data Analysis
Anastasia Basharina-Freshville
(University College London)

## Slide 2
The Waveform Fitter
Reader:
  Automatically calculates the number of points per event (this is a setting on the digitiser and can vary for runs)
  Creates a tree with a histogram for each waveform
Fitter:
  Fit the first part of the wavelength with a “pol0 + gaus” fit to obtain the baseline
Baseline
Note: waveform example taken from file 079 from July 2015

## Slide 3
  Carry out a baseline subtraction on the waveform using p0 = baseline
  Fit the baseline subtracted waveform in three parts/regions:
The Waveform Fitter (2)
Part 1:
Fit the rise time with a Gaussian
(3 parameters)
Part 2:
Fit a Gaussian around the peak
(3 parameters)
Part 3:
Fit the decay time with an exponential
(2 parameters)

## Slide 4
  Combine step 3 into one fit, with each part of the fit corresponding to a defined region of the waveform:
    8 free parameters
    4 fixed parameters to define regions
  Take the integral of the waveform histogram for a defined gate length (x-axis) starting from a position along the x-axis defined by the digitiser settings during each run
    Cut on a Χ2/NDF value to remove pile up from the “integral” histogram
The Waveform Fitter (3)
This region could use improvement, but for now the Χ2/NDF for single vs double/pile up waveforms is a good way to identify pile up.
Time, ns
Amplitude, mV

## Slide 5
Single vs Pile Up Waveforms
Example waveforms taken from run 053, August 2016:
  1.98 mm collimator, no absorbers
  300 – 500 kHz / 0.2 – 0.27 nA scattering foil current rate
Χ2/NDF looks to be a good way to distinguish between single events and pile up (checked for many waveforms!)
See slide 19 for resulting zoomed in Χ2/NDF distribution for run 053
Single waveform:
Double waveform:
Χ2/NDF = 23
Χ2/NDF = 2922

## Slide 6
Tuesday 2nd of August 2016

## Slide 7
Run 003: Rate < 25 kHz, Wheel Position 0/1.4 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 122881 | 6344 ± 1.0 | 2.819 ± 0.042
< 33 | 66973 | 6339 ± 0.9 | 2.582 ± 0.035
ADC equivalent: run 001 |  |  | 
None | 1901507 | 6360 ± 0.1 | 2.219 ± 0.008

## Slide 8
Run 013: Rate < 25 kHz, Wheel Position 100/18.9 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 92161 | 4116 ± 1.1 | 5.596 ± 0.074
< 33 | 65094 | 4114 ± 1.4 | 5.574 ± 0.096
ADC equivalent: run 009 |  |  | 
None | 2146863 | 3851 ± 0.1 | 5.127 ± 0.010
Note: There was a “beam drop” in between the ADC run and the waveform run during which the ADC mean shifted

## Slide 9
Run 031: Rate - 20 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 104974 | 6524 ± 0.9 | 2.105 ± 0.038
< 33 | 53473 | 6522 ± 0.7 | 2.079 ± 0.030
ADC equivalent: run 029 |  |  | 
None | 1963124 | 6628 ± 0.1 | 2.115 ± 0.007
x 103

## Slide 10
Run 034: Rate - 25 kHz, 7.2 mm PMMA Plate (placed in same position as wheel), equivalent to wheel position 40/8.4 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 95293 | 5286 ± 1.0 | 2.862 ± 0.042
< 33 | 61349 | 5283 ± 1.3 | 2.938 ± 0.058
ADC equivalent: run 032 |  |  | 
None | 2023737 | 5405 ± 0.1 | 3.037 ± 0.008

## Slide 11
Run 036: Rate - 12 kHz, 7.2 mm PMMA Plate (placed upstream of the beam), equivalent to wheel position 40/8.4 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 96997 | 5372 ± 0.9 | 2.869 ± 0.045
< 33 | 62426 | 5371 ± 1.2 | 2.933 ± 0.058
ADC equivalent: run 035 |  |  | 
None | 1953740 | 5484 ± 0.2 | 2.922 ± 0.008

## Slide 12
Run 038: Rate - 20 kHz, 13.4 mm PMMA Plate (placed in same position as wheel), equivalent to wheel position 80/15.4 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 97729 | 4240 ± 1.0 | 4.196 ± 0.054
< 33 | 66985 | 4237 ± 1.2 | 4.180 ± 0.068
ADC equivalent: run 037 |  |  | 
None | 2111164 | 4344 ± 0.1 | 4.397 ± 0.009
Note: There was a “beam drop” during the waveform acquisition, use event number < 50000

## Slide 13
Run 040: Rate - 7 kHz, 13.4 mm PMMA Plate (placed upstream of the beam), equivalent to wheel position 80/15.4 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 50000 | 4810 ± 1.4 | 3.836 ± 0.073
< 33 | 32606 | 4810 ± 1.8 | 3.742 ± 0.091
ADC equivalent: run 039 |  |  | 
None | 1615642 | 4911 ± 0.2 | 3.927 ± 0.010

## Slide 14
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 95632 | 6558 ± 0.7 | 2.318 ± 0.030
< 33 | 48059 | 6556 ± 0.8 | 1.997 ± 0.028
ADC equivalent: run 041 |  |  | 
None | 1814614 | 6644 ± 0.1 | 2.101 ± 0.007
Run 042: Rate - 20 kHz, No Absorbers

## Slide 15
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 96600 | 6439 ± 1.4 | 2.555 ± 0.062
< 33 | 51406 | 6437 ± 1.6 | 2.247 ± 0.058
ADC equivalent: run 043 |  |  | 
None | 2327671 | 6517 ± 0.1 | 2.067 ± 0.007
Run 044: Rate - 50 kHz, No Absorbers
Note: Beam drops to 0 during waveform run acquisition. This is also reflected in the ADC spectrum, but is zoomed here.

## Slide 16
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 141313 | 6320 ± 0.7 | 2.227 ± 0.030
< 33 | 78347 | 6301 ± 0.4 | 2.228 ± 0.025
ADC equivalent: run 045 |  |  | 
None | 2777435 | 6406 ± 0.1 | 2.003 ± 0.007
Run 046 Rate - 85 kHz, No Absorbers
x 103

## Slide 17
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 141313 | 6291 ± 0.6 | 2.488 ± 0.025
< 33 | 76064 | 6292 ± 0.8 | 2.289 ± 0.028
ADC equivalent: run 047 |  |  | 
None | 4491099 | 6373 ± 0.1 | 1.963 ± 0.006
Run 048: Rate - 140 kHz, No Absorbers

## Slide 18
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 184321 | 6254 ± 0.6 | 2.246 ± 0.022
< 33 | 91010 | 6252 ± 0.7 | 2.455 ± 0.027
ADC equivalent: run 049 |  |  | 
None | 5654307 | 6335 ± 0.1 | 1.917 ± 0.006
Run 050: Rate - 250 kHz, No Absorbers

## Slide 19
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 273409 | Unable to fit spectrum | 
< 33 | 118449 | Unable to fit spectrum | 
ADC equivalent: run 049 |  |  | 
None | 2011560 | Unable to fit spectrum | 
Run 053: Rate - 500 kHz / 0.2 – 0.27 nA Scattering Foil Current, No Absorbers

## Slide 20
Validation of the Χ2/NDF Method for Run 053 Using Visual Scanning (1)
Method:
  Look at every 100th event and if the Χ2/NDF is < 33 then save the waveform histogram
  1,224 out of 273,409 events saved and scanned by eye
  0.9 % of events were not perfect, but in the majority of cases the second peak is small and outside of the integration gate:
  99.1 % of events were found to be “good”:

## Slide 21
Validation of the Χ2/NDF Method for Run 053 Using Visual Scanning (2)
The resulting Χ2/NDF distribution:
The distribution is very similar to that of Slide 19 even with such low statistics
  Is this indication of a gain shifting with rate problem?
Rebinned

## Slide 22
Wednesday 3rd of August 2016

## Slide 23
Run 055: Rate – 8 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 96376 | 6689 ± 0.6 | 2.495 ± 0.031
< 33 | 41124 | 6676 ± 0.7 | 1.607 ± 0.032
ADC equivalent: run 054 |  |  | 
None | 1330743 | 6767 ± 0.2 | 2.111 ± 0.008

## Slide 24
Run 057: Rate – 170 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 129025 | 6185 ± 0.5 | 2.807 ± 0.028
< 33 | 69310 | 6183 ± 0.7 | 2.391 ± 0.031
ADC equivalent: run 056 |  |  | 
None | 2133455 | 6249 ± 0.2 | 1.987 ± 0.008

## Slide 25
Run 059: Rate – 125 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 95233 | 6201 ± 0.6 | 2.745 ± 0.032
< 33 | 52553 | 6202 ± 0.8 | 2.361 ± 0.033
ADC equivalent: run 058 |  |  | 
None | 2160452 | 6276 ± 0.2 | 2.146 ± 0.009

## Slide 26
Run 061: Rate – 250 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 178177 | 6157 ± 0.7 | 2.092 ± 0.030
< 33 | 87775 | 6141 ± 0.4 | 2.554 ± 0.029
ADC equivalent: run 060 |  |  | 
None | 3705478 | 6218 ± 0.2 | 1.849 ± 0.009

## Slide 27
Run 063: Rate – 300 - 500 kHz / 0.2 nA Scattering Foil Current, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 184321 | 6158 ± 0.7 | 2.137 ± 0.031
< 33 | 76969 | 6157 ± 0.8 | 2.651 ± 0.035
ADC equivalent: run 062 |  |  | 
None | 5180357 | 6210 ± 0.2 | 1.597 ± 0.012

## Slide 28
Run 065: Rate – 0.35 nA Scattering Foil Current, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 187393 | 4460 ± 1.2 | 10.588 ± 0.105
< 33 | 87150 | 4435 ± 2.1 | 11.821 ± 0.174
ADC equivalent: run 064 |  |  | 
None | 2774077 | 4741 ± 0.5 | 8.896 ± 0.039

## Slide 29
Run 067: Rate – 230 kHz, 0.55 Ø Collimator, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 141313 | 6086 ± 1.7 | 1.866 ± 0.125
< 33 | 78335 | 6103 ± 1.5 | 2.715 ± 0.077
ADC equivalent: run 066 |  |  | 
None | 3730863 | 6153 ± 0.3 | 2.169 ± 0.019

## Slide 30
Run 069: Rate – 250 kHz, 0.55 Ø Collimator, 8 mm Off-axis, Left of Centre, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 138241 | 6086 ± 1.1 | 2.110 ± 0.065
< 33 | 71453 | 6078 ± 0.6 | 2.497 ± 0.057
ADC equivalent: run 068 |  |  | 
None | 4107566 | 6154 ± 0.1 | 2.160 ± 0.014

## Slide 31
Run 071: Rate – 250 kHz, 0.55 Ø Collimator, 8 mm Off-axis, Top of Centre, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 141313 | 6107 ± 1.0 | 2.119 ± 0.054
< 33 | 75536 | 6110 ± 1.1 | 2.578 ± 0.052
ADC equivalent: run 070 |  |  | 
None | 38851999 | 6171 ± 0.2 | 2.557 ± 0.012

## Slide 32
Run 073: Rate – 250 kHz, 0.55 Ø Collimator, 8 mm Off-axis, Bottom of Centre, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 153601 | 6096 ± 0.9 | 1.875 ± 0.041
< 33 | 84435 | 6096 ± 1.0 | 2.500 ± 0.052
ADC equivalent: run 072 |  |  | 
None | 4052432 | 6159 ± 0.2 | 2.366 ± 0.011

## Slide 33
Run 075: Rate – 110 kHz, 0.55 Ø Collimator, 16 mm Off-axis, Bottom of Centre, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 98305 | 2335 ± 0.8 | 11.495 ± 0.077
< 33 | 89590 | 2332 ± 0.9 | 11.352 ± 0.078
ADC equivalent: run 074 |  |  | 
None | 4639029 | 2431 ± 0.1 | 9.417 ± 0.018

## Slide 34
Run 078: Rate – 100 kHz, 0.55 Ø Collimator, 16 mm Off-axis, Left of Centre, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 159745 | 2349 ± 1.1 | 10.624 ± 0.139
< 33 | 147988 | 2350 ± 1.1 | 10.823 ± 0.135
ADC equivalent: run 077 |  |  | 
None | 4862886 | 2469 ± 0.1 | 9.309 ± 0.018

## Slide 35
Hexagonal Module + 8” PMT (with –VE Base)

## Slide 36
Run 104: Rate – 170 kHz, Brass Absorbers to Fully Stop Protons: Neutron and Gamma Spectrum

## Slide 37
Run 106: Rate – 18 kHz, Brass Absorbers to Fully Stop Protons: Neutron and Gamma Spectrum

## Slide 38
Run 109: Rate – 130 kHz, 1.98 Ø Collimator, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 98305 | Unable to fit spectrum | 
> 100 & event number > 50000 | 28390 | 15968 ± 5 | 5.580 ± 0.068
ADC equivalent: run 108 |  |  | 
None | 3056216 | 17835 ± 0.2 | 2.239 ± 0.040
Note: ADC mean appears to have been shifting during the waveform acquisition, therefore it’s difficult to get a spectrum to fit