## Slide 1: Clatterbridge 24th – 25th November 2016 ADC Data Analysis
Anastasia Basharina-Freshville
(University College London)

## Slide 2
Summary of Measurements Made
Optical module setup used:
  2” Hamamatsu R13089-100-11 PMT with negative HV UCL soldered base
  3 cm x 3 cm x 5 cm cuboid ENVINET standard scintillator
  Coupled with BC-630 Saint Gobain silicone optical gel (refractive index = 1.465) http://www.crystals.saint-gobain.com/uploadedFiles/SG-Crystals/Documents/Organic%20Product%20Accessories%20Data%20Sheet.pdf
Thursday the 24th August 2016:
  All measurements carried out at -900 V
  Measurements of beam operating at different rates (30 kHz – 400 kHz)
  Monitoring of correlation between the rate measured on the oscilloscope and the displayed scattering foil current
  Waveform data collected with the LeCroy oscilloscope in order to more accurately measure rates
  Waveform data collected with the digitiser and the LeCroy oscilloscope simultaneously using a 50 Ω LEMO splitter and barrels
  Peak current estimate at -900 V (for an amplitude of 447 mV and a width of ~16 ns, for a threshold of -100 mV): 11.2 mA

## Slide 3
Summary of Measurements Made
Wednesday the 3rd August 2016:
  All measurements carried out at -900 V
  Energy-range calibration scan with the Clatterbridge calibration wheel carried out at low rates starting at 50 kHz for a gate of 150 ns
  Measurements with two PMMA plate thicknesses corresponding to those of the wheel to establish whether we go through one or more of the wheel segments with the beam

## Slide 4
Energy Resolution and Reproducibility of Measuremets: 60 MeV Beam
1.98 mm Ø collimator
Energy resolution: ~2.5 %
5965
6329
6412
Time difference between runs:
004: 12.27
013: 14.17
026: 15.38

## Slide 5
Reproducibility of Measurements: Wheel Position 0
1.98 mm Ø collimator
Wheel position 0 (1.4 mm water equivalent, 1.2 mm PMMA)
Energy resolution: 2.7 – 2.8 %
Time difference between runs:
001: 09.58
030: 11.27

## Slide 6
Comparison of Wheel Position 40 and PMMA
1.98 mm Ø collimator
Wheel position 40 (7.2 mm PMMA), PMMA plate used: 7.24 mm
Energy resolution: 2.5 – 3.5 %

## Slide 7
Comparison of Wheel Position 80 and PMMA
1.98 mm Ø collimator
Wheel position 80 (13.3 mm PMMA), PMMA plate used: 13.4 mm
Energy resolution: 4.9 – 5.3 %
σ is similar for measurements with the wheel and the PMMA plates
  We can conclude that we are therefore only going through one wheel segment.

## Slide 8
Energy-Range Calibration
Measurements at higher water equivalent depth do not produce a spectrum that can be fit with the Lan-Gaus convolution
μ measured using a simple Gaussian around the peak for all points for consistency
Note: The μ returned by Lan-Gaus and by the simple Gaussian fits are usually within ± 10 ADC counts of each other
A more sensible looking calibration than August 2016 data?
Bragg Peak Line

## Slide 9
Energy-Range Calibration (Zoom Around Bragg Peak Points)
Bragg Peak Line

## Slide 10
Energy-Range Calibration from Simulations
Simulations carried out in GEANT v10.0.2 (SL 5)
No beam line, scintillation turned off
MeV to ADC calibration taken from MeV vs ADC plots on Slide 14 (data taken during this test beam)
One point has lower ADC than expected. Also seen in August 2016 data and simulations.
Bragg Peak Line

## Slide 11
Energy-Range Calibration from Simulations (Zoom Around Bragg Peak Points)
Bragg Peak Line

## Slide 12
Energy Resolution as a Function of Energy (FWHM)
NOTE:
Systematic errors still need to be estimated for this data set
“Guesstimate” for now just for the fit, with ± 1% for the lowest energy points
A great χ2/NDF even with incorrect errors

## Slide 13
Energy Resolution as a Function of Energy (σ)
NOTE:
Systematic errors still need to be estimated for this data set
“Guesstimate” for now just for the fit, with ± 1% for the lowest energy points
A great χ2/NDF even with incorrect errors

## Slide 14
Linearity/Energy Scan: -900 V
1st degree polynomial fit over range shown on plot
p0 gives an estimate of quenching in the scintillator (cf 20.49 MeV from simulations for EJ PVT scintillator):
  18.06
Previously, the lowest proton beam energy we tested to was ~ 32 MeV
Harder to fit the mean at lower energies due to lower statistics
But even at lower energies we still look to be linear
Previous PMMA plates used not below this line

## Slide 15
Rate Tests
1.98 mm Ø collimator
As rate increases, ADC mean increases and energy resolution appears to improve.
The ADC mean increasing is the opposite trend to the data we collected in August 2016 with the Hamamatsu active divider base!

## Slide 16
Scope Rate – Scattering Foil Current Calibration
The plot currently shows scope rate measured at Clatterbridge using the “duration” function
For complete accuracy we need to analyse scope waveform data to get accurate rates
Once we have an accurate measurement of this calibration we should be able to rely on the scattering foil current for rate measurements

## Slide 17
Radiation Damage Assessment
207Bi tests carried out at -1700 V:
  before test beam: 21/11/2016
  after test beam: to be done
Total radiation dose received by 2” OM (not yet accounting for collimators etc.): 428.4 monitor units (0.9 Gy/monitor unit) in 367 minutes of running.
Scaling for 1.98 mm collimator diameter from a 17 mm beam diameter and converting to Gy: 1.5 Gy
No noticeable worsening in resulting energy resolution.
Date | HV Supply | DAQ | ΔE/E (%)
21/11/16 | UCL | UCL | 9.02 ± 0.22
Clatterbridge Test Beam: 24-25/11/16 |  |  | 
16/01/16 | UCL | UCL | 7.98 ± 0.15
This scintillator has now received a total of 4.2 Gy (August 2016) + 1.5 Gy (November 2016): 5.7 Gy