Range uncertainties with QuARC

---
Uncertainty sources
- fit: 0.02 - 0.05 mm (because of 1% light uncertainty in sheet)
- Stack WET:
  - 0.045 mm (error on mean of PKF measurements)
  - Energy dependence of WET: 0.07 – 0.15 mm (30 – 130 MeV, stopping power ratio)
- HIT reference ranges: 0.04 mm (Parodi 2012)
- Stack position: 0.1 mm (10 cm air)
- Stack angle: 0.3 mm (4 degrees rotation plausible because of laser and markers)
- Birks’ constant: 0.1 - 0.2 mm (b = 0.07 – 0.09 mm/MeV)
- Combined uncertainty: 0.4 mm
  + add 0.2mm for each PMMA block used as a degrader (possible rotation)

---
WET measurement uncertainty
- Problem: the WET measurement uncertainty seems to be underestimated
- Offsets in Delta R: Uncertainty on PMMA degrader thickness
- Slope in each Delta R section: Uncertainty on stack WET
Offset
(degrader WET)
Slope
(stack WET)

---
WET measurement uncertainty
- Possible explanation: Positioning uncertainty (rotation)
- The positioning of the scintillator stack at an angle to the beam will lead to an increased observed thickness of the stack.
- 4 degrees misalignment lead to a thickness increase of 0.3mm over 127mm stack length.
- 3 degrees is a realistic scenario given the wobbly table and potential uncertainty of the laser and the reference lines drawn on top of the case.
d’ = d/sin(α)
d
beam
α
d’

---
PTW Peakfinder uncertainty
If stack WET is 127.5 mm instead of 127.1 mm:

---
PTW Peakfinder uncertainty
If PMMA degraders are also corrected (and also offset in air):
air: +0.23mm, PMMA1: -0.03mm, PMMA2: +0.2mm, PMMA3: +0.3mm

---
Birks’ constant
- Birks’ constant is determined from a fit to a “most trusted” curve at 90mm range (maximum range in stack, no degrader used)
- However, the fit is very sensitive to small shape variations, error bar size and peak position in stack.
- Birks’ constant could be different in every sheet due to ageing
b = 0.092 mm/MeV
b = 0.0706 mm/MeV

---
Stopping power ratio
- WET depends on beam energy! Because stopping power ratio depends on beam energy (it decreases with increasing energy)
- That means the WET of the stack seen by a low energy beam is shorter than the WET measured with a high energy beam

---
Stopping power ratio
How much shorter?
~0.07mm for 50MeV beam and 0.15mm for 140MeV beam
SPR data for polystyrene and water from NIST database (PSTAR)
Average SPR for range meas (140MeV)
Average SPR for stack WET meas

---
Conclusion
- WET is energy-dependent!
- Total uncertainty on range is larger than expected
- Really important to develop an accurate positioning system for detector