## Slide 1
Birks’ constant measurement
Laurent Kelleter

## Slide 2: Birks constant measurement
MedAustron
HIT Nov18
HIT Apr19
Increase of Birks constant with decreasing beam energy.
Question: Why?

## Slide 3: Birks constant fit
Birks is fitted and a large value is returned
Birks is fixed at 0.09 mm/MeV. Fit good except for first sheets
Answer: Fit is pulled up by weird data points
Next question: What causes the calibration to fail in the first few sheets?

## Slide 4: Answer: Protons hit sensor!
HIT nov18
MedAustron
White pixels from protons that hit sensor directly
Spot size 29.8 mm FWHM
E = 62.4 MeV
Spot size 21.5 mm FWHM
E = 59.15 MeV
All images are increased in brightness to make it easier to see the white pixels by direct proton hits. The images are NOT saturated at the Bragg peak!

## Slide 5: Different energies/spot sizes at HIT
HIT nov18
Spot size 32.4 mm FWHM
E = 48.47 MeV
HIT nov18
Spot size 13.8 mm FWHM
E = 116.39 MeV
Protons hit sensor
No protons hit sensor

## Slide 6: Similar spot sizes HIT/MA
MedAustron
E = 62.4 MeV
Spot size 21.5 mm FWHM
HIT nov18
Spot size 20.2 mm FWHM
E = 78.87 MeV
Protons in sensor go further because of higher energy
Also protons in sensor, but travel less far
=> Beam position in sensor is approximately the same for HIT and MA

## Slide 7: Conclusion
Calibration of PDL curves fails because protons enter the sensors directly
This is due to a larger beam spot size at HIT compared with MedAustron at the same proton energy
The direct hits increase the light output in the first sheets
Which reduces the peak-to-plateau ratio
And ultimately increases the fit value of Birks constant
This effect gets stronger with decreasing beam energy because of the increasing beam spot size