## Slide 1: Clatterbridge 2nd – 3rd August 2016 ADC Data Analysis
Anastasia Basharina-Freshville
(University College London)

## Slide 2
Summary of Measurements Made
Optical module setup used:
  2” Hamamatsu R13089-100-11 PMT with negative HV active divider base (made by Hamamatsu)
  3 cm x 3 cm x 5 cm cuboid ENVINET standard scintillator
  Coupled with BC-630 Saint Gobain silicone optical gel (refractive index = 1.465) http://www.crystals.saint-gobain.com/uploadedFiles/SG-Crystals/Documents/Organic%20Product%20Accessories%20Data%20Sheet.pdf
Tuesday the 2nd August 2016:
  All measurements carried out at -900 V
  Energy-range calibration scan with the Clatterbridge calibration wheel carried out at rates starting at 25 kHz for a gate of 150 ns
  Measurements with two PMMA plate thicknesses corresponding to those of the wheel to establish whether we go through one or more of the wheel segments with the beam
  Measurements with changing beam parameters in order to increase the rates measured
  Lots of unsaturated waveform data collected with the digitiser and the new LeCroy oscilloscope
  Peak current estimate at -900 V (for an amplitude of 421 mV and a width of ~29.2 ns, for a threshold of -100 mV): 12 mA

## Slide 3
Summary of Measurements Made
Wednesday the 3rd August 2016:
  All measurements carried out at -900 V
  A repeat of the rate tests, pushing the beam up to operation with the same parameters as used for treatment (scattering foil current: 0.001 – 1.8 nA)
  Lots of unsaturated waveform data collected with the digitiser
  Beam uniformity measurements with 8 mm and 16 mm off-axis 0.5 mm Ø collimators
  A repeat of the energy-range calibration scan with the Clatterbridge calibration wheel using finer steps
  8” PMT + Hexagonal OM measurements:
    Neutron and ϒ measurements
    Background (beam off) measurement
    Energy resolution measurement

## Slide 4
Energy Resolution and Reproducibility of Measuremets: 60 MeV Beam
1.98 mm Ø collimator
Energy resolution: ~2.1 %
PMT gain drift?
6619
6635
6758
0.24 %
difference
Time difference between runs taken on the same day:
029: 13.47
041: 15.37

## Slide 5
Reproducibility of Measurements: Wheel Position 0
1.98 mm Ø collimator
Wheel position 0 (1.4 mm water equivalent, 1.2 mm PMMA)
Energy resolution: 2.2 - 2.6 %
PMT gain drift?
6360
6238
6443
7616
Time difference between runs taken on the same day:
001: 10.55
027: 12.17
1.3 %
difference

## Slide 6
Comparison of Wheel Position 40 and PMMA
1.98 mm Ø collimator
Wheel position 40 (7.2 mm PMMA), PMMA plate used: 7.24 mm
Energy resolution: ~3 %
Can we conclude that we are therefore only going through one wheel segment?

## Slide 7
Comparison of Wheel Position 80 and PMMA
1.98 mm Ø collimator
Wheel position 80 (13.3 mm PMMA), PMMA plate used: 13.4 mm
Energy resolution: 3.9 – 4.4 %
σ increases by ~5% and the μ shifts in both directions for the PMMA plate measurements
  “Wheel” σ < “PMMA plate” σ, therefore are we going through one wheel segment?
  Why the change in σ and the shift in μ with the PMMA plate?

## Slide 8
Energy-Range Calibration
Measurements at higher water equivalent depth do not produce a spectrum that can be fit with the Lan-Gaus convolution
μ measured using a simple Gaussian around the peak for all points for consistency
Note: No noticeable difference when using Lan-Gaus μ values
Bragg Peak Line

## Slide 9
Energy-Range Calibration
Measurements at higher water equivalent depth do not produce a spectrum that can be fit with the Lan-Gaus convolution
μ measured using a simple Gaussian around the peak for all points for consistency
Note: No noticeable difference when using Lan-Gaus μ values
Gradient changes between these two points. Why?
Bragg Peak Line
Data from the two different days doesn’t seem to follow the same “pattern”?

## Slide 10
Energy-Range Calibration (Zoom Around Bragg Peak Points)
Bragg Peak Line

## Slide 11
Energy-Range Calibration from Simulations
Simulations carried out in GEANT v10.0.2 (SL 5)
No beam line, scintillation turned off
MeV to ADC calibration taken from MeV vs ADC plots on Slide 13 (data taken during this test beam)
Two points have lower ADC than expected

## Slide 12
Energy-Range Calibration from Simulations (Zoom Around Bragg Peak Points)
Bragg Peak Line

## Slide 13
Energy Resolution as a Function of Energy
NOTE:
Systematic errors still need to be estimated for this data set
“Guesstimate” for now just for the fit, with ± 1% for the lowest energy points
Reasonable looking χ2/NDF even with incorrect errors

## Slide 14
Linearity/Energy Scan: -900 V
1st degree polynomial fit over range shown on plot
p0 gives an estimate of quenching in the scintillator (cf 20.49 MeV from simulations for EJ PVT scintillator):
  11.36 from 02.08.16
  19.37 from 03.08.16
Previously, the lowest proton beam energy we tested to was ~ 32 MeV
Harder to fit the mean at lower energies due to lower statistics
Looks like we could be linear when looking at the higher energy points
Previous PMMA plates used not below this line

## Slide 15
Rate Tests: 02.08.2016
1.98 mm Ø collimator
As rate increases, ADC mean decreases and energy resolution improves (but probably within our systematic error).

## Slide 16
Rate Tests: 03.08.2016
1.98 mm Ø collimator
Not as much of a noticeable change in μ with increase in rates as for tests on previous day except for the first increase in rate.
Visually, as rate increases, σ increases. However, from the fit as rate increases, σ decreases (but Χ2/NDF gets worse) and therefore energy resolution (from calculation) improves.
Operating with beam parameters the same as for treatment!
Waveform data for these and higher rates to be analysed.

## Slide 17
Beam Uniformity Measurements
0.5 mm Ø collimator
The beam appears to be uniform 8 mm away from the centre.
16 mm away from the centre is 1mm away from the beam edge, where other factors and effects come into play.
In the future it would be good to have collimators made with more various off-axis distances and repeat the test.
For now we can conclude that the beam is uniform and that reducing rates using a collimator whilst keeping nominal beam operating conditions is a valid method to get the beam down to measurable rates.

## Slide 18
Background Measurement
1.98 mm Ø collimator
Beam Off!

## Slide 19
Radiation Damage Assessment
207Bi tests carried out at -1700 V:
  before test beam: 27-29/07/2016
  after test beam: 04/08/2016
Total radiation dose received by 2” OM (not yet accounting for collimators etc.): 1361.4 monitor units (0.9 Gy/monitor unit) in 508.35 minutes of running.
No noticeable difference in resulting energy resolution.
Date | HV Supply | DAQ | ΔE/E (%)
27/07/16 | Portable | UCL | 7.42 ± 0.07
28/07/16 | UCL | UCL | 7.38 ± 0.08
29/07/16 | UCL | Digitiser | 7.57 ± 0.29
Clatterbridge Test Beam: 02-03/08/16 |  |  | 
04/08/16 | UCL | UCL | 7.23 ± 0.07

## Slide 20
8” OM: Neutron and Gamma Measurements
Measurements carried out with the 8” PMT + Hexagonal Scintillator OM at -900 V
Brass collimators and other bits around the nozzle to stop all protons
Rate: ~18 kHz

## Slide 21
8” OM: Background Measurement
Measurement carried out with the 8” PMT + Hexagonal Scintillator OM at -900 V
Beam off!

## Slide 22
8” OM: Energy Resolution Measurement
1.98 mm Ø collimator
No absorbers (60 MeV beam)
Rate: 130 kHz
ΔE/E: 2.24 ± 0.04 %
Last ΔE/E measurement, carried out in July 2015 at – 800 V with a 2 mm Ø collimator at a rate of 50 Hz: 2.18 ± 0.03 %
No degradation in measurement