## Slide 1: Clatterbridge 24th – 25th November 2016 Waveform Data Analysis
Anastasia Basharina-Freshville
(University College London)

## Slide 2
Thursday 24th of November 2016:
Rate Tests

## Slide 3
Run 003: Rate - 30 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 196609 | 6231 ± 0.3 | 2.718 ± 0.015
< 60 | 179404 | 6234 ± 0.3 | 2.715 ± 0.015
ADC equivalent: run 001 |  |  | 
None | 32768 | 6141 ± 0.1 | 2.544 ± 0.07
Lowest rate file
Use to train Χ2 cut
Use a cut of Χ2 < 60 for following runs

## Slide 4
Run 006: Rate - 50 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 199681 | 6441 ± 0.4 | 2.755 ± 0.015
< 60 | 177618 | 6442 ± 0.4 | 2.731 ± 0.014
ADC equivalent: run 004 |  |  | 
None |  | 6345 ± 0.1 | 2.504 ± 0.006

## Slide 5
Run 009: Rate - 100 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 199681 | 7291 ± 0.4 | 2.267 ± 0.014
< 60 | 145458 | 7289 ± 0.4 | 2.035 ± 0.014
ADC equivalent: run 007 |  |  | 
None |  | 7155 ± 0.1 | 2.779 ± 0.006

## Slide 6
Run 012: Rate - 222 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 184321 | 12828 ± 1.5 | 3.901 ± 0.036
< 60 | 10564 | Unable to fit due to low statistics | 
ADC equivalent: run 010 |  |  | 
None |  | 13015 ± 0.2 | 1.853 ± 0.005

## Slide 7
For a proton beam rate of ~ 200 kHz the pulse height of the signal drastically increases, resulting in a higher mean of the ADC distribution
Applying a cut of Χ2 < 60 reduces the mean of the ADC distribution from ~ 13000 to ~ 7000, which is around the mean we see at rates < 200 kHz
When looking at the ADC mean as a function of event number without applying a Χ2 cut (Slide 6) some fluctuation is seen
Looking at the ADC mean as a function of event number for Χ2 < 60 (on this slide) shows that only events with an ADC mean of ~ 7000 are left
Statistics are too low to fit the ADC spectrum with a Χ2 < 60, but this is an interesting observation
Similar observation in other 200 kHz runs, see later.

## Slide 8
Run 015: Rate - 50 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 150529 | 6536 ± 0.4 | 2.660 ± 0.016
< 60 | 130476 | 6538 ± 0.4 | 2.616 ± 0.018
ADC equivalent: run 013 |  |  | 
None |  | Unable to fit spectrum due to shifting during run | 

## Slide 9
Run 018: Rate - 300 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 242689 | 6789 ± 2.0 | 10.925 ± 0.120
< 60 | 10738 | 6709 ± 2.6 | 16.641± 0.082
ADC equivalent: run 016 |  |  | 
None |  | 7155 ± 0.1 | 2.779 ± 0.006
At higher rates we start to see the same “beam shifting” pattern as with the Hamamatsu active divider base
Is this an indication that there’s a problem with the PMT itself?

## Slide 10
Run 020: Rate - 200 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 159745 | Unable to fit due to double peak | 
< 60 | 12424 | Unable to fit due to low statistics | 
ADC equivalent: run 019 |  |  | 
None |  | 12291 ± 0.5 | 4.006 ± 0.010

## Slide 11
For a proton beam rate of ~ 200 kHz the pulse height of the signal drastically increases, resulting in a higher mean of the ADC distribution
Applying a cut of Χ2 < 60 reduces the mean of the ADC distribution from ~ 13000 to ~ 7000, which is around the mean we see at rates < 200 kHz
When looking at the ADC mean as a function of event number without applying a Χ2 cut (Slide 10) some fluctuation is seen
Looking at the ADC mean as a function of event number for Χ2 < 60 (on this slide) shows that only events with an ADC mean of ~ 7000 are left
Statistics are too low to fit the ADC spectrum with a Χ2 < 60, but this is an interesting observation
Similar observation in other 200 kHz runs, see earlier and later.
Why the double peak? According to my notes the scattering foil current started to drop towards the end of the run. Subsequent ADC mode spectra also showed beam shifting.

## Slide 12
Run 025: Rate - 400 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, % | 
None | 196609 | 4346 ± 3.9 | 8.632 ± 0.425 | 
< 60 | 90931 | 4500 ± 3.2 | 21.50 ± 0.16 | 
ADC equivalent: run 024 |  |  |  | 
None |  | 4676 ± 0.6 |  | 11.912 ± 0.037
Making a Χ2 < 60 cut produces a nicer beam shape, but it is still very wide
The spread in the mean of the ADC is smaller compared to August 2016 data, but we can still see the mean shifting

## Slide 13
Run 027: Rate - 50 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, % | 
None | 150529 | 6138 ± 0.4 | 3.002 ± 0.014 | 
< 60 | 124760 | 6137 ± 0.4 | 3.018 ± 0.015 | 
ADC equivalent: run 026 |  |  |  | 
None |  | 5981 ± 0.1 |  | 2.552 ± 0.005

## Slide 14
Thursday 24th of November 2016:
Scope and Digitiser 50 Ω LEMO Splitter Tests

## Slide 15
Run 029: Rate - 50 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, % | 
None | 156673 | 2914 ± 0.4 | 4.712 ± 0.029 | 
< 10 | 149259 | 2916 ± 0.4 | 4.768 ± 0.029 | 
ADC equivalent: run 028 |  |  |  | 
None |  | 3046 ± 0.1 |  | 3.125 ± 0.008
Our signal is now split in half – we expect the Χ2 distribution to look different, so look at lowest rate file to choose the Χ2 cut value
Use Χ2 < 10
We also expect to see ~ half of the ADC mean value we previously saw, as we do

## Slide 16
Run 031: Rate - 200 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 119809 | 5599 ± 0.7 | 4.482 ± 0.032
< 10 | 13235 | Unable to fit due to low statistics | 
ADC equivalent: run 030 |  |  | 
None |  | 5786 ± 0.1 | 2.815 ± 0.007

## Slide 17
For a proton beam rate of ~ 200 kHz the pulse height of the signal drastically increases, resulting in a higher mean of the ADC distribution
Applying a cut of Χ2 < 10 reduces the mean of the ADC distribution from ~ 5500 to ~ 3000, which is around half of the mean we see at rates < 200 kHz when using a splitter
When looking at the ADC mean as a function of event number without applying a Χ2 cut (Slide 16) some fluctuation is seen
Looking at the ADC mean as a function of event number for Χ2 < 10 (on this slide) shows that only events with an ADC mean of ~ 3000 are left
Statistics are too low to fit the ADC spectrum with a Χ2 < 10, but this is an interesting observation
Similar observation in other 200 kHz runs, see earlier.

## Slide 18
Run 033: Rate - 400 kHz, No Absorbers
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 147457 | Unable to fit spectrum | 
< 10 | 63152 | 2196 ± 1.2 | 14.064 ± 0.162
ADC equivalent: run 032 and 034 |  |  | 
None | Unable to fit spectrum due to shifting during running |  | 

## Slide 19
Friday 25th of November 2016:
Wheel Energy-Range Calibration Tests

## Slide 20
Run 002: Rate - 50 kHz, Wheel Position 0/1.4 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 92161 | 5818 ± 0.8 | 3.584 ± 0.032
< 60 | 64268 | 5830 ± 0.8 | 3,146 ± 0.041
ADC equivalent: run 001 |  |  | 
None |  | 5672 ± 0.1 | 2.702 ± 0.007

## Slide 21
Run 008: Rate - 50 kHz, Wheel Position 100/18.9 mm WE
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 95233 | 3234 ± 1.0 | 8.766 ± 0.059
< 60 | 83300 | 3235 ± 1.0 | 7.865 ± 0.014
ADC equivalent: run 007 |  |  | 
None |  | 3139 ± 0.1 | 6.247 ± 0.012

## Slide 22
Friday 25th of November 2016:
Wheel Equivalent PMMA Plates Tests

## Slide 23
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 92161 | 4714 ± 0.6 | 4.204 ± 0.035
< 60 | 68154 | 4718 ± 0.7 | 4.067 ± 0.042
ADC equivalent: run 031 |  |  | 
None |  | 4609 ± 0.1 | 4.156 ± 0.008
Run 032: Rate - 50 kHz, 7.2 mm PMMA Plate (placed in same position as wheel), equivalent to wheel position 40/8.4 mm WE

## Slide 24
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 95662 | 4614 ± 0.5 | 3.442 ± 0.035
< 60 | 72025 | 4617 ± 0.5 | 3.341 ± 0.038
ADC equivalent: run 033 |  |  | 
None |  | 4495 ± 0.1 | 3.478 ± 0.009
Run 034: Rate - 50 kHz, 7.2 mm PMMA Plate (placed upstream of the beam), equivalent to wheel position 40/8.4 mm WE

## Slide 25
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 98305 | 3679 ± 0.7 | 5.451 ± 0.045
< 60 | 81365 | 3684 ± 0.7 | 4.997 ± 0.057
ADC equivalent: run 035 |  |  | 
None |  | 3570 ± 0.1 | 5.258 ± 0.010
Run 036: Rate - 50 kHz, 13.4 mm PMMA Plate (placed in same position as wheel), equivalent to wheel position 80/15.4 mm WE

## Slide 26
Χ2/NDF cut | Number of events | ADC mean, ADC counts | ΔE/E FWHM, %
None | 95233 | 3653 ± 0.6 | 4.376 ± 0.059
< 60 | 80121 | 3657 ± 0.5 | 4.087 ± 0.055
ADC equivalent: run 037 |  |  | 
None |  | 3539 ± 0.1 | 4.930 ± 0.011
Run 038: Rate - 50 kHz, 13.4 mm PMMA Plate (placed upstream of the beam), equivalent to wheel position 80/15.4 mm WE

## Slide 27
Conclusions
Although to somewhat a lesser extent we still see the “beam shifting” problem when running at higher rates
In order to investigate this:
  Use the detector emulator to test the detector at various rates in a controlled environment
  Whilst getting to know the emulator we could pulse an LED at different rates and see whether we see any effects from this