## Slide 1: Proton Calorimetry
Simon Jolly
University College London

## Slide 2: SuperNEMO
Neutrinoless double beta decay detector using NEMO3’s tracker-calorimeter technique Target sensitivity: T½ > 1026 years → <mν> <0.04 – 0.1 eV
Modular detector with a planar geometry
1 module (of 20) consists of:
Source foil:
  5 kg (total of 100 kg) of 40 mg/cm2 (4 x 2.7 m2)
  82Se (high Qββ, long T1/22vββ, proven enrichment technology): starting baseline
  150Nd and 48Ca being considered depending on enrichment possibilities
Tracker: ~2000 drift cells in Geiger mode → particle identification (for background suppression)
Calorimeter: ~550 scintillator blocks + PMTs → energy and time of flight measurements of particles
Passive shielding surrounding each module
2 m (assembled, ~0.5 m between source and calorimeter)
6 m
4 m

## Slide 3: SuperNEMO Calorimeter Development
SuperNEMO calorimeter consists of 550 Optical Modules (wrapped scintillator block + PMT).
Needed to achieve improved resolution:
Scintillator needed to be organic plastic:
  High light yield.
  Low electron back-scattering.
  High radiopurity.
  Fast timing.
Extensive R&D programme resulted in final design:
  ElJen EJ-200 polyvinyltoluene (PVT) scintillator.
  Hamamatsu 8” PMT (32% QE at 400 nm).
  Hexagonal scintillator block directly coupled to hemispherical PMT face.
  Teflon + Mylar wrapping.

## Slide 4: Optimised Optical Module
Wrapping:
Sides: 75 μm of PTFE (Teflon) ribbon
Sides and entrance face: 12 μm of Mylar
EJ-200 hexagonal PVT block:
276 mm diameter
193 mm deep, minimum thickness between PMT and scintillator: 100 mm
R5912-MOD Hamamatsu 8” PMT:
Maximum quoted QE: 33%
32% QE at 400 nm
Wavelength (nm)
QE (%)

## Slide 5: So What…?
With this fantastic energy resolution of 7.5% FWHM at 1 MeV can we use a SuperNEMO Optical Module for PBT energy measurement?
Very high intensity of events at a proton beam (10’s MHz):
  A proton beam delivers a random number of protons per bucket, which will worsen the energy resolution measured.
  We require 1 proton per bucket for a good detector response.
  But time constant of scintillator (1 ns) makes this possible.
Scintillator quenching for protons:
  For a plastic scintillator, the scintillator response is nonlinear with the amount of energy deposited in it.
  Amount of deviation → “quenching”.
  Characterised by Birk’s law:
  Becomes important for large dE/dx and ionisation density → important for protons, which have a large dE/dx when they slow down.
Energy range:
  SuperNEMO optimised for electrons from 0.5 – 4 MeV for double beta decay.
  For proton therapy we require ~O(100 MeV).
dY/dx light yield per unit path length
dE/dx energy lost by particle per unit path length
kB relates density of ionisation to energy loss
= 0.207 mm/MeV
S absolute scintillation efficiency

## Slide 6: Simulated Stopping Distance
Simulations of SuperNEMO scintillator vs Water Equivalent:
PVT is “water equivalent” for stopping distance and spread, as is PS.
One to one conversion for water phantoms.
Is this important to radiotherapy physics…?
Proton Beam Energy, MeV | Mean stopping distance, SCINT (mm) | Mean stopping distance, WATER (mm) | σ stopping distance, SCINT (mm) | σ stopping distance, WATER (mm)
60 | 30.21 | 30.54 | 0.33 | 0.33
200 | 255.4 | 257.1 | 2.48 | 2.44
300 | 505.9 | 509.9 | 4.64 | 4.78

## Slide 7: Proton Stopping Power
Treatment plans constructed using X-ray CT.
Electron density is NOT the same as proton stopping power.
Hounsfield unit conversion factor introduces range uncertainty.

## Slide 8: Range Uncertainty (T. Lomax)
Tumour shrinkage after 5 weeks

## Slide 9: Proton CT (N. Allinson, PRaVDA)
Use proton imaging to reduce range uncertainties in treatment planning:
  Measure proton entry and exit positions and trajectories with trackers.
  Measure energy with Residual Energy Resolving Detector (range telescope or calorimeter).
  Calculate proton most likely path to reconstruct image.

## Slide 10: Proton CT Image Reconstruction

## Slide 11: Proton Calorimetry
For Proton CT, need to read out proton energy quickly and accurately:
  Energy resolution <1% 𝜎.
  Rate 1–10 MHz for reasonable imaging time.
Detector requirements overlap with range QA:
  Daily measurements are faster but less comprehensive: really just want to verify that the energy of the beam specified by the control system has the correct range in water (Water Equivalent Path Length).
  Morning range QA normally verifies a few energies at known depths:
    Proton counting in water-equivalent phantom.
    Setup and measurement can be time consuming…
  Requirements for improved QA system:
    Better than 1% σ resolution across all energies.
    Easier system setup and faster data acquisition and readout.
See if SuperNEMO calorimeter works for protons…

## Slide 12: Clatterbridge Cancer Centre
62 MeV Scanditronix cyclotron provides 60 MeV protons (31 mm in water) to treatment room through double scattering.
Beam time provided for research.
We’ve had 2-day shifts every few months.
Already made interesting observations with our equipment about the treatment beam…
Need much lower proton fluence for our measurements than clinical settings.
Rate reduction achieved through:
  Various collimators (0.5–10 mm)
  Ion source gas supply.
  Ion source discharge current.
  Cyclotron sector focussing.
  RF phasing (wouldn’t recommend it…).

## Slide 13: Equipment Setup
CAEN DT5751 Digitiser:
Dual-gate signal integration
→ On-the-fly pulse shape analysis
→ Neutron/gamma discrimination

## Slide 14: Experimental Tests

## Slide 15: Results: Fitted Data
Mirror Landau Tail
Fitting function:
Convolution of Gaussian
and mirror Landau + Landau on the right
→ ΔE/E: 1.58 ± 0.27 % FWHM (for ~40 MeV due to quenching!)
Compared to 1.48 % FWHM from simulations

## Slide 16: High Rate Tests: Pulse Pile-Up
Protons scattering off the collimator increases with diameter decrease
More than 1 proton per bucket for larger diameter collimators, causing peak to widen
To avoid pile up:
Optimal collimator: 2mm diameter
Keep integrating window small to only collect single proton pulse

## Slide 17: A Smaller, Faster Detector
We have already achieved the target energy resolution: 0.7% σ with
The next step is to do this for very high rates of 1–10 MHz with a compact design:
Reduce the size of the PMT and the scintillator to improve timing and make the design nozzle-mountable.
Negative HV PMT base to remove decoupling capacitor (not fast enough discharge).
2” Hamamatsu R13089-100-11 PMT with negative HV active divider base
  3 cm x 3 cm x 5 cm cuboid ENVINET/NUVIA PolyStyrene standard scintillator
  Coupled with BC-630 Saint Gobain silicone optical gel
Wrapped in 75 μm of PTFE (Teflon) ribbon on the sides and 12 μm of Mylar on the sides and entrance face

## Slide 18: Small Module Results
→ σ/E: 0.89 ± 0.11 % σ
at 25 kHz
We are still at our < 1 % σ target!

## Slide 19: Resolution: Energy Dependence
Energy of protons incident on scintillator varied by placing absorbers (PMMA plates and calibration wheel) of known thickness ~1.8 m upstream of the optical module.
dependence!

## Slide 20: Resolution: Linearity
We want to run the PMT at higher voltages (can run at up to -1500V) as this will increase the PMT’s collection efficiency and will improve the energy resolution.
BUT we have a LOT of light (tens of thousands of photo-electrons) so we need to make sure we are not saturating the PMT.
Look at linearity:
For -900V:
Deviation from linearity < 2%!

## Slide 21: Beam Test Conclusions
What have we learned?
The scintillator performs just as well for single protons as it does for electrons!
Making the module smaller does what it’s supposed to:
  Improves timing (good measurements up to around 300kHz, compared to 1 kHz for original 8” module),
  No detrimental effect on resolution.
But…
  We still can’t handle rates approaching 1 MHz.
  Despite Hamamatsu’s promises to the contrary, we think the PMTs have a frequency-dependent gain.
Interesting discoveries about Clatterbridge beam:
  Nonlinear time distribution of protons (bunches of bunches...).
  Close to nozzle edge, energy falls off.
  Building complete simulation to compare to Clatterbridge/UCL measurements:

## Slide 22: MedAustron 250 MeV Tests
Having tested detector at 60 MeV, needed to make high energy test to examine performance and test linearity:
  Longer, 40 cm scintillator block to absorb 250 MeV protons.
  Same PMT and readout.
Through OMA, MedAustron offered us 2 nights of beam tests:
  62 MeV up to 252 MeV.
  Able to drop rate down with chopper adjustment: ran 1 kHz–1 MHz.
  Custom collimators to reduce intensity and beam size.

## Slide 23: MedAustron Setup

## Slide 24: MedAustron Results

## Slide 25: So What…?
Our goal was originally to develop a calorimeter to act as the energy measurement stage for a proton CT system:
  Needed better than 1% resolution and rates in the region of 1–10 MHz.
  Managed to achieve the resolution; rates limited by electronics.
  Work will continue: discussions with PRaVDA and Loma Linde.
Clinical steer to provide fast energy/range QA tool to work at clinical rate.
Needs a change in design philosophy: also take advantage of water equivalence of plastic scintillator.
Treatment Nozzle

## Slide 26: Segmented Calorimeter
PVT and PS are both helpfully water equivalent.
Segment block into slices and read out light from each slice individually.
Integrate signal from many protons: very large output from 1010/s.
Minimum slice width will depend on manufacture: aiming for < 2 mm.
Use photodiodes for readout: poor light detectors but stable and cheap with large dynamic range.
Resolution set by slice width and variation in scintillator light output.
Protons

## Slide 27: Segmented Calorimeter Design
Laurent Kelleter has built preliminary model in Geant4:
  2 mm slices of plastic scintillator with mylar wrapping.
  Currently integrating photodiode readout.
STFC IPS grant application currently pending approval: working with NUVIA a.s. in Czech Republic to produce our scintillator sheets: manufacturing challenging!
Need to characterise light quenching to reconstruct Bragg curve: pencil beams only.
Fit to measured curve drastically improves mean range measurement: do you need range or range spread…?
Dose deposited
Light emitted
dY/dx | light yield per unit path length
dE/dx | energy lost by particle per unit path length
kB | relates density of ionisation to energy loss = 0.207 mm/MeV
S | absolute scintillation efficiency

## Slide 28: Fast Treatment Plan Verification
Take segmented calorimeter: add 2D tracking to front face.
Still nozzle-mounted and self-contained.
Read out X/Y profile and integrated range of individual pencil beams.
Detector read out fast enough to match minimum spot dwell time (3–20 ms).
Fast reconstruction of water-equivalent treatment plan.
Protons
Tracking layer

## Slide 29: TERA: Proton Range Radiography
Don’t need to prove the principle using scintillator sheets: TERA have done it for us!
Proton Range Radiography:
  Gas Electron Multiplier (GEM) tracking.
  2 mm PVT scintillator sheets fibre coupled to Silicon PhotoMultipliers.
Can’t use this exact setup:
  Designed for single protons for pCT.
  They get “good enough” proton range by looking at end-of-range only.
  SiPMs expensive, high gain devices: not appropriate for high light output with full beam intensity.

## Slide 30: Future Plans
Continue development of single calorimeter module for proton
CT and lower rate applications:
  Well characterised.
  The fewer channels the better: single block also means more light per proton per detector.
Work on segmented calorimeter design to produce water equivalent path length detector:
  Resolution better than 2 mm: much better with appropriate fit.
  “Immediate” readout (a few seconds).
  Need >150 sheets for 32 cm: start with 20 sheets and do fast measurement at Clatterbridge.
Full design aims to be gantry mounted: can characterise multiple fields.
Fast treatment plan verification very promising, but needs work to get segmented calorimeter working before adding tracking:
  Tracking and range measurement need to be fast enough to read out data with suitable resolution within spot dwell time.
  Needs electronics to synchronise.