## Slide 1: Using a mixed He/C beam for monitoring intrafractional motion
Lennart Volz1,2 * & Laurent Kelleter3 *, R. Radogna3, C. Graeff4, C. Schömers5, S. Brons5, S. Jolly3 and J. Seco1,2
PTCOG 58
1 Department of Biomedical Physics in Radiation Oncology, DKFZ, Hedielberg, GER
2 Department of Physics and Astronomy, Heidelberg University, Heidelberg, GER
3 Department of Physics and Technology, University College London, London, UK
4 Gesellschaft für Schwerionenforschung GmBH, Darmstadt, GER
5 Heidelberg Ion-Beam Therapy Center, Heideberg, GER
* Equal contribution

## Slide 2: Ion-beam radiotherapy: range uncertainty
Target
Organ at risk
Ideal: Prescribed dose to target, minium dose to healthy tissue
Needed: Online treatment monitoring
Tumor
Organ at Risk
Patient
Nozzle

## Slide 3: He/C mixed beam: See what you treat
1 Graeff et al. (2018); 2Mazzuconi et al. (2018); Lühr et al. (2011) (http://dEdx.au.dk)
4He2+ and 12C6+ range at the same velocity 3
RBE optimized He/C SOBP1 log scale!
Patient
Detector
4He2+ and 12C6+ have approx. the same A/q1,2
 Simulatenous acceleration to same velocity
Helium range ~3 times Carbon range
10% Helium less than 0.5% target RBE dose 2

## Slide 4: Experiments at Heidelberg Ion-Beam Therapy Center
Experiments at HIT QA cave
Separate Helium and Carbon beams
Novel range telescope developed at UCL
  12cm stack of thin 2-3mm scintillation sheets
  Read out by flat panel CMOS sensor
Geant4 simulations of the setup
UCL range telescope prototype at HIT
Mylar foil
Polystyrene sheets
CMOS sensor
Poster at PHY04
Nozzle
Detector

## Slide 5: Simple PMMA geometries: Setup
PMMA degrader (WET ≈ 21.9 cm), different air gaps after ~5cm build up
12C: 219.19 MeV/u (~10cm range), 4He: 220.5 MeV/u (~30cm range), beam FWHM ~8mm
Experiment: Intensity 8x107 particles/s
Experiment
Simulation

## Slide 6: Simple PMMA geometries: Data processing
Experiment
- Preliminary
Quenched scintillation light output
Adding the He and C signal together (10:1 ratio)
Simulation

## Slide 7: Simple PMMA geometries: How sensitive is the method?
- Preliminary
Experiment
Simulation
2mm wide → 22% of beam

## Slide 8: Simple PMMA geometries: online range monitoring
- Preliminary
PMMA degrader (WET ≈ 21.9 cm)
2mm PMMA slab covering half the field after ~5cm build up
3x3cm2 scanned field, ~3x106 p/spot, E=210MeV/u, 8x107 p/s, ~10 mm FWHM
5cm
2mm
12C peak
~9.1cm
Beam

## Slide 9: Anthropomorphic phantom: What could we use this method for?
1 Niebuhr et al. (2019); 2 Wieser et al. (2017)
In-house developed ADAM deformable pelvis phantom 1
Designed for multimodal imaging QA
Carbon treatment plan with MatRad 2
  Use central energyE=324.26MeV/u
Intensity 4He: 7x108, 12C: 8x107
Beam FWHM 4He: 7mm, 12C: 8mm
Inflation of rectum with rectal balloon
Axial
Sagittal

## Slide 10: Anthropomorphic phantom: What could we use this method for?
- Preliminary
Beams eye view
Bladder
Prostate
Rectum + balloon

## Slide 11: Anthropomorphic phantom: What could we use this method for?
- Preliminary
Beams eye view
Bladder
Prostate
Rectum + balloon

## Slide 12: Anthropomorphic phantom: What could we use this method for?
- Preliminary
Beams eye view
Bladder
Prostate
Rectum + balloon

## Slide 13
6/4/2019 |
Take home message
A mixed Helium/Carbon beam offers the potential of online treatment monitoring and post-treatment QA
First experimental results further confirm previous proof-of-concept investigations
Many challenges in proving the feasibility of acceleration…
Next steps:
Currently under investigation!
Which patients would benefit from the technique?
How to use the additional information?

## Slide 14: Backup

## Slide 15: Backup: Production of a mixed beam
Extraction?
Extract 4He+/12C3+
Pre-acceleration?
Not possible at HIT LINAC optimized for
Would in general be possible
Synchrotron?
Stable acceleration, extraction, scanning, dose rate etc..?
More investigation needed!
Schematic of the HIT accelerator complex

## Slide 16: Backup: Methane source spectrum
- Courtesy of Tim Winkelmann (HIT)
Contamination from previous run with CO2

## Slide 17: Backup: ADAM PETer rotation
2cm PMMA degrader
Beam E=324.26

## Slide 18: Backup: Geant4 simulation setup
Separate Helium and Carbon beams
Geant4.10.2.p03
G4HadronPhysics_QGSP_BIC_HP, G4IonBinaryCascade, G4EmStandardPhysics_option4, G4StoppingPhysics
Birks constant kB=0.126 MeV/mm
Nozzle ~2.03mm WET (G4_WATER), Nozzle-isocenter distance 1.02m
PMMA degrader density optimized to match measured degrader RSPs
CT scan loaded from HU to Density and HU to Material lookup table taken directly from CT scanner QA