## Slide 1: QuARC: A Quality Assurance Range Calorimeter for Proton Therapy
Sonia Escribano-Rodriguez1, Saad Shaikh1, Raffaella Radogna2, Simon Jolly1
1University College London, Physics & Astronomy, London, United Kingdom
2University of Bari, Physics, Bari, Italy
DEPARTMENT OF PHYSICS AND ASTRONOMY
Queen Mary University of London
Wednesday 5th April 2023

## Slide 2
Overview
Proton Therapy
Challenges in PT
QuARC
Analytical model
Experimental results
Clinical application
Geant4 Simulation
Future work

## Slide 3
Proton therapy
1
Proton Therapy (PT) is a modality of radiotherapy that utilises proton beams with energies between 70 MeV and 250 MeV to irradiate cancerous structures.
(F Aliyah et al 2021 IOP Conf. Ser.: Earth Environ. Sci. 927 012042)

## Slide 4
Proton therapy
PT offers a highly localised depth-dose distribution, reducing the damage to healthy tissues nearby. Specially indicated for:
(Johns Hopkins Medicine)
Tumours close to critical structures
Depth tumours
2
Paediatric tumours

## Slide 5
Proton therapy
Better dose conformation due to the finite energy loss of the protons. The energy deposition is characterised by the Bethe-Bloch formula
3

## Slide 6
Proton therapy
Better dose conformation due to the finite energy loss of the protons. The energy deposition is characterised by the Bethe-Bloch formula
3

## Slide 7
Proton therapy
(R. Orecchia, A. Zurlo, A. Loasses et al. 1996. Eur. J. Cancer, volume 34)
3
Better dose conformation due to the finite energy loss of the protons. The energy deposition is characterised by the Bethe-Bloch formula

## Slide 8
Proton therapy
(R. Orecchia, A. Zurlo, A. Loasses et al. 1996. Eur. J. Cancer, volume 34)
3
Better dose conformation due to the finite energy loss of the protons. The energy deposition is characterised by the Bethe-Bloch formula

## Slide 9
Proton therapy
To cover the whole tumoral region a spread-out Bragg peak (SOBP) is created by combining beams with different energies.
(Geng C, Gates D, Bronk L, Ma D and Guan F. 2018 Phys Med.51:13-21)
4

## Slide 10
FLASH Proton therapy
FLASH radiotherapy is a novel technique that offers ultra-high dose-rate. While the chemical and biological mechanisms of the FLASH effect are undefined, early preclinical evidence suggests FLASH radiotherapy offers additional healthy tissue sparing effects with complete dose delivery in less than 100 ms.
For FLASH proton therapy this corresponds to average currents of at least 600 nA to the patient, compared to current clinical standard of 1–10 nA [1].
5

## Slide 11
Proton therapy: challenges
Knowing precisely the proton range is a technical challenge, and the range uncertainties prevent taking full advantage of PT benefits.
To ensure accurate dose delivery, daily QA measurements are required. The properties are usually measured with a commercial ionisation-chamber.
- Time-consuming (up to 1h) → compromise between speed and accuracy.
- Current detectors used for QA cannot cope with the requirements of FLASH PT.
6

## Slide 12
QuARC: Quality Assurance Range Calorimeter
Clinical devices are time consuming and can’t cope with the requirements for FLASH proton therapy, which exposes the need for a device capable of working at FLASH and regular rates in a quicker way.
Develop a detector to make fast, accurate measurements of proton range
7

## Slide 13
QuARC
Optically-isolated polystyrene scintillator sheets of 105 mm x 105 mm x 3 mm.
Segmented design: each sheet is coupled to a photodiode to measure light output, so each layer is individually sampled.
8

## Slide 14
QuARC
Photodiodes coupled to modular ADC converters, read-out by an FPGA.
The scintillator sheets are housed in a custom plastic 3D-printed holder in groups of 32 sheets.
9

## Slide 15
QuARC
The scintillator modules are placed in a light-tight case, with two mylar windows for beam entrance.
10

## Slide 16
QuARC set up
11

## Slide 17
QuARC model
The light output of the scintillator is subject to quenching effects → non-linear response in the scintillation light output to absorbed dose.
Birks’ law for light quenching
The aim is to develop a model that describes the depth-light curve in a quenching scintillator and allow the extraction of information about the beam range.
12

## Slide 18
QuARC model
Analytical model deployed to fit depth-light curve obtained by the PD to reconstruct Bragg curve and recover proton range in real-time [2]. The model combines Bortfeld’s description of the Bragg curve with Birks’ empirical law describing scintillator light quenching as a function of LET.
13

## Slide 19
QuARC model
4-fit-parameter model [2].
14

## Slide 20
UCLH experiment
3 modules tested
Beam energies 70-150 MeV
15
White & clear sheets
Clear sheets only

## Slide 21
UCLH experiment
Calibration is required with a shoot-through measurement to acknowledge the differences in light output between the sheets (130 MeV).
16
White & clear sheets
Clear sheets only

## Slide 22
UCLH experiment
Background measurements for comparison and future subtraction.
17

## Slide 23
UCLH experiment
Initial calibrated data for a 90 MeV beam
18
Clear sheets only
White sheets only

## Slide 24
UCLH experiment II
  Black: Sheet Average Light Output
  Blue: Fitted Quenched Depth-Light Curve
  Green: Reconstructed Depth-Dose Curve
  Magenta: UCLH Reference Dose Curve
130 MeV
Accurate to 0.1 mm
19

## Slide 25
UCLH experiment: conclusions
Proof of principle tested
Mylar foil gives better results than painting the sheets
Range reconstruction accurate to 0.1 mm
Total time ~10 mins (including setup time)
Working detector at clinical rates
Higher beam energies?
FLASH dose rate?
19

## Slide 26
20
The Christie experiment
Testing 4 modules (32 x 3mm sheets each) with clinical and FLASH currents.
Beam energies between 70 and 245 MeV.

## Slide 27
Initial calibration is required with a shoot-through measurement.
245 MeV, front
245 MeV, back
The Christie experiment
21

## Slide 28
Comparison of the integrated photodiode charge for clinical and FLASH currents for a 245 MeV proton beam. The normalised calibration ratio shows that the light output is within 1% of each other.
22
The Christie experiment

## Slide 29
FLASH measurements at 245 MeV: initial data.
The Christie experiment
23

## Slide 30
Christie Beam Test: Results
24
245 MeV

## Slide 31
Christie Beam Test: Results
24
245 MeV
150 MeV

## Slide 32
Christie Beam Test: Results
24
150 MeV
245 MeV

## Slide 33
The Christie experiment: conclusions
Difference in light output due to productions mechanisms
Detector working at energies up to 245 MeV
Detector can handle FLASH rates
Clinical application?
25

## Slide 34
Clinical application
26
A compact version of the QuARC has been designed for the Clatterbridge proton therapy centre in Liverpool, UK, as part of an upgrade to their clinical control system.
The detector will have a compact computing solution on-board to both perform data-acquisition and fitting.
Clinical devices provide the software to obtain the range value without the need for any coding/analysis during the measurement → GUI.

## Slide 35
Web GUI
View live fit results in web from any device.
27
No software required.

## Slide 36
Geant4 Simulation
Compare simulated results with experimental data.
Investigate Birks’ constant values.
29

## Slide 37
Future work
Further development of web-based GUI to allow data display and detector control without any specialised software on the user-end.
Improve scintillator composition and construction to optimise light output.
Deployment of fitting routines onto an FPGA to increase fitting speed.
Geant4 simulation to benchmark experimental results and constrain Birks’ constant.
Additional beam tests to further characterise detector performance with clinical and FLASH beams.
30

## Slide 38
Acknowledgements

## Slide 39
Thank you!
[1] S. Jolly et al. “Technical Challenges for FLASH Proton Therapy”. In Physica Medica 78 (2021). DOI: 10.1016/j.ejmp.2020.08.005
[2] L. Kelleter and S. Jolly. “A Mathematical Expression for Depth-Light Curves of Therapeutic Proton Beams in a Quenching Scintillator”. In: Medical Physics 47.5 (Feb. 2020), pp. 2300–2308. DOI: 10.1002/mp.14099