# DOCTOR FRY’S DOCTORED AND FRIED LEAD

# VISUAL INSPECTION LAB: FAULTS IDENTIFIED

Version: 2011-10-20

| LEAD NO. | 13A MAINS PLUG |  |  |  |  | CABLE | IN-LINE CONNECTOR | IEC CONNECTOR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CASE | PINS | FUSE | WIRES | CLAMP |  |  |  |
| 1 |  | Wobbly Neutral pin | Replaced with bolt | Loose copper strands | Clamp on 3 wires not tightened |  |  | Missing strain relief + cut insulation |
| 2 | Cracked – fixed with Araldite |  |  |  |  | Nick in insulation near IEC connector | 3-way IEC socket to 2-way plug (no earth) | Missing screw on strain relief |
|  |  |  |  |  |  |  |  | Cracked case |
| 3 | Rubber hood cut to fit cable | Unshrouded Live + Neutral (1 of 2) | Corroded | Copper strands fitted with ferrules |  |  |  |  |
|  |  | Corroded (1 of 3) |  |  |  |  |  |  |
| 4 |  |  | Replaced with paper clip | Copper strands tinned | Missing |  | 2-way “lead connector” wrong way around | Class II strain relief broken/cut |
| 5 | Melted + burned |  |  |  |  | Squashed (trolley ran over it) |  | Earth socket broken off |
|  |  |  |  |  |  |  |  | Live and Neutral reversed |
|  |  |  |  |  |  |  |  | Missing cable clamp |
| 6 |  | Unshrouded Live + Neutral (2 of 2) |  |  | Replaced with twisted wire | USA “wrong” colour L = Black N = White E = Green |  |  |
|  |  | Corroded (2 of 3) |  |  |  | Insulation shaved |  |  |
| 7 |  | Neutral pin sheared off | Fitted with 50A car fuse |  | Broken | Melted with flame |  | L, N + E exposed by slicing of plastic end |
| 8 | Part of lid missing exposing Live pin |  | Blown fuse with wire soldered to it |  |  | 2-core cable used for IEC with earth |  | Knotted with cable clamp missing |
|  |  |  |  |  |  | Loudspeaker wire |  |  |
| 9 | Fixing bolt does not secure lid | Plastic earth pin |  | Live + Neutral reversed | Not tightened |  | In-line switch |  |
| 10 |  | Screw head sheared off | Replaced with 30A fuse wire | Wrong colours for L + N (Red + Black) |  | “Permanent” wiring cable |  | Earth not connected |

| LEAD NO. | 13A MAINS PLUG |  |  |  |  | CABLE | IN-LINE CONNECTOR | IEC CONNECTOR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CASE | PINS | FUSE | WIRES | CLAMP |  |  |  |
| 11 |  | Plastic shroud cut off Live pin |  | Earth wire “hanging on” by a few strands | Cable not clamped |  | 3-way “lead connector” socket to 2-way plug with E to cable clamp | Melted |
|  |  | No screw on Live |  |  |  |  |  |  |
| 12 |  |  |  | Earth wire cut off |  | USA coding: L = Black N = White E = Green “wrong” colour for UK | IEC plug connected to 13A plug + Twinbloc conversion |  |
|  |  |  |  | Insulation stripped too much |  |  | Two cables joined with uninsulated terminal block |  |
| 13 | Top of lid sawn off and sellotaped on | Corroded (3 of 3) |  | Wires loose Screw terminals not tight |  | Badly repaired by twisting + sellotape |  |  |
|  |  |  |  |  |  | Knotted |  |  |
| 14 | Base nearly sawn in half | Fuse holder rivet sheared off Live pin |  | Insulation not stripped off |  | Melted with soldering iron | No cable clamps on 3-way “lead connectors” |  |
| 15 |  |  | Replaced with bolt |  |  | Chewed through by a rat (really!) | 3-way “lead connector” wrong way around |  |
|  |  |  |  |  |  |  | L and E reversed |  |
|  |  |  |  |  |  |  | Missing cable clamp screws |  |
| 16 |  |  |  | Live connected to Earth pin | Outer insulation should be clamped | Insulation damaged |  |  |
|  |  |  |  | Nick in Neutral insulation |  |  |  |  |
|  |  |  |  | Copper strands tinned |  |  |  |  |
| 17 | Attempt to re-use a non-rewireable plug |  | Metal bolt touching live fuse | Earth wire soldered with only 2 strands | Cable clamp cut | Melted near IEC connector |  |  |
|  |  |  |  | L and N reversed |  |  |  |  |
| 18 | US Mains plug for 110V equipment ! | Earth pin sawn off to fit into UK adaptor |  |  |  | Melted near IEC connector | Black L to Blue N White N to Brown L |  |
|  |  | UK 2-pin adaptor cut |  |  |  |  |  |  |
| 19 | “Schuko” plug with earth into UK shaver adaptor (no earth) |  | Replaced with bolt |  |  | Knotted |  | L, N + E exposed by slicing of plastic end |
|  |  |  |  |  |  | Crushed near IEC connector |  |  |
| 20 | Rubber hood cut to fit cable | Corroded (3 of 3) | Replaced with paper clip |  |  |  |  | Missing strain relief |
|  | Biro top used to hold cover on |  |  |  |  |  |  |  |

| LEAD NO. | 13A MAINS PLUG |  |  |  |  | CABLE | IN-LINE CONNECTOR | IEC CONNECTOR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CASE | PINS | FUSE | WIRES | CLAMP |  |  |  |
| 21 |  |  | Replaced with bolt | L & N Reversed |  | Melted and crushed | 2-pin shaver adapter | Cut in strain relief |
| 22 |  | Unshrouded L & N |  |  |  | Perished rubber red & black insulation |  | L & N reversed |
|  |  |  |  |  |  |  |  | Wire cable clamp |
| 23 | Cut |  | No fuse | Earth detached |  |  | Twisted & sellotaped | No L & N guard |
|  |  |  |  |  |  |  |  | Sockets exposed |
| 24 |  |  | Brass Rod | Wrong type of wire No earth | Ineffective |  | Wrong gender | Melted cable |
| 25 |  |  | No fuse | All wires wrong |  | Saw cut | Terminal block | Exploded ! |
|  |  |  |  |  |  |  | Exposed wires |  |
| 26 |  |  | Wire shorting fuse | L & N reversed | No clamp | Cut |  | Cut wire |
| 27 |  |  |  |  |  |  |  |  |
| 28 |  |  |  |  |  |  |  |  |
| 29 |  |  |  |  |  |  |  |  |
| 30 |  |  |  |  |  |  |  |  |