|  | MCQ Safety of Medical Electrical Equipment |
| --- | --- |
| For MSc Students | MSc in Physics and Engineering Medicine |
| Department | Medical Physics and Bioengineering |
| Module Name | Clinical Practice |
| Module Portico Code | MPHYGB17 |
| Lecturer | Dr Martin Fry |
|  | This exercise has been given to provide experience in using the standard BS EN 60601-1:2006, and does not form part of the assessment of this module. |

INSTRUCTIONS:

To complete this MCQ you will need to consult the standard BS EN 60601-1:2006 Medical Electrical Equipment Part 1: General requirements for basic safety and essential performance.

A UCL-licensed pdf copy of this document can be downloaded from the British Standards Online website: https://bsol.bsigroup.com/en/BsolHomePage/ using a UCL computer and UCL account.

Please do not print this document as it is nearly 400 pages long! Think of the financial and environmental costs.

In order to complete this MCQ you will in fact find it easier to search electronically using appropriately chosen keywords.

Choose the most appropriate answers from the selection given. For some questions there may be more than one correct answer.

Write the letter(s) you select in the empty box under the word ”Answer”.

You may add “keywords”, “clause, table or figure” numbers and “comments” in the other empty boxes to provide “evidence” of the validity of your answer (particularly useful if there is any dispute regarding the correct answers!)

You should also enter your letter selections in the summary of answers table at the beginning of this document.

Summary of answers for MCQ:

Safety of Medical Electrical Equipment:

| Question | Answer(s) |  | Question | Answer(s) |
| --- | --- | --- | --- | --- |
| 1 |  |  | 26 |  |
| 2 |  |  | 27 |  |
| 3 |  |  | 28 |  |
| 4 |  |  | 29 |  |
| 5 |  |  | 30 |  |
| 6 |  |  | 31 |  |
| 7 |  |  | 32 |  |
| 8 |  |  | 33 |  |
| 9 |  |  | 34 |  |
| 10 |  |  | 35 |  |
| 11 |  |  | 36 |  |
| 12 |  |  | 37 |  |
| 13 |  |  | 38 |  |
| 14 |  |  | 39 |  |
| 15 |  |  | 40 |  |
| 16 |  |  | 41 |  |
| 17 |  |  | 42 |  |
| 18 |  |  | 43 |  |
| 19 |  |  | 44 |  |
| 20 |  |  | 45 |  |
| 21 |  |  | 46 |  |
| 22 |  |  | 47 |  |
| 23 |  |  | 48 |  |
| 24 |  |  | 49 |  |
| 25 |  |  | 50 |  |

| 1. | ### APPLIED PART: |  |  |  |
| --- | --- | --- | --- | --- |
|  | According to BS EN 60601-1:2006, the APPLIED PART is a part of ME EQUIPMENT that in NORMAL USE: |  |  |  |
|  | a | Necessarily comes into physical contact with the PATIENT for ME EQUIPMENT or an ME SYSTEM to perform its function. |  |  |
|  | b | Can be brought into contact with the PATIENT or OPERATOR. |  |  |
|  | c | Needs to be touched by the PATIENT or OPERATOR. |  |  |
|  | d | Must not be touched by the OPERATOR |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 2. | ### MAINS PART: |  |  |  |
| --- | --- | --- | --- | --- |
|  | The definition of the MAINS PART includes the following: |  |  |  |
|  | a | Electrical circuit that is intended to be connected to the SUPPLY MAINS. |  |  |
|  | b | The MAINS PART includes all conductive parts that are not separated from the SUPPLY MAINS by at least one MEANS OF PROTECTION. |  |  |
|  | c | The PROTECTIVE EARTH CONDUCTOR is regarded as a part of the MAINS PART. |  |  |
|  | d | The PROTECTIVE EARTH CONDUCTOR is not regarded as a part of the MAINS PART. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 3. | ### CATEGORY APG EQUIPMENT: |  |  |  |
| --- | --- | --- | --- | --- |
|  | A rating for ME EQUIPMENT or an ME EQUIPMENT part complying with specified requirements on construction, marking and documentation in order to avoid sources of ignition in a FLAMMABLE ANAESTHETIC MIXTURE with: |  |  |  |
|  | a | Air |  |  |
|  | b | Oxygen |  |  |
|  | c | Nitrogen |  |  |
|  | d | Nitrous oxide |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 4. | ### CLASS I EQUIPMENT: |  |  |  |
| --- | --- | --- | --- | --- |
|  | EQUIPMENT in which protection against electric shock does not rely on BASIC INSULATION only, but also relies on: |  |  |  |
|  | a | A protective earth conductor |  |  |
|  | b | A functional earth conductor |  |  |
|  | c | DOUBLE INSULATION |  |  |
|  | d | REINFORCED INSULATION |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 5. | ### CLASS II EQUIPMENT: |  |  |  |
| --- | --- | --- | --- | --- |
|  | EQUIPMENT in which protection against electric shock does not rely on BASIC INSULATION only, but also relies on: |  |  |  |
|  | a | A protective earth conductor |  |  |
|  | b | A functional earth conductor |  |  |
|  | c | DOUBLE INSULATION |  |  |
|  | d | REINFORCED INSULATION |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 6. | ### TYPE B APPLIED PART: |  |  |  |
| --- | --- | --- | --- | --- |
|  | APPLIED PART complying with the specified requirements of this standard to provide protection against electric shock, particularly regarding: |  |  |  |
|  | a | Allowable PATIENT LEAKAGE CURRENT |  |  |
|  | b | Allowable PATIENT AUXILIARY CURRENT |  |  |
|  | c | Insulation method |  |  |
|  | d | Not suitable for DIRECT CARDIAC APPLICATION |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 7. | ### TYPE BF APPLIED PART: |  |  |  |
| --- | --- | --- | --- | --- |
|  | APPLIED PART which: |  |  |  |
|  | a | Provides a higher degree of protection against electric shock than TYPE B APPLIED PARTS |  |  |
|  | b | Is an F-TYPE APPLIED PART |  |  |
|  | c | Insulation method |  |  |
|  | d | Not suitable for DIRECT CARDIAC APPLICATION |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 8. | ### TYPE CF APPLIED PART: |  |  |  |
| --- | --- | --- | --- | --- |
|  | APPLIED PART which: |  |  |  |
|  | a | Provides a higher degree of protection against electric shock than TYPE BF APPLIED PARTS |  |  |
|  | b | Is an F-TYPE APPLIED PART |  |  |
|  | c | Insulation method |  |  |
|  | d | Is intended for direct cardiac connection only |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 9. | ### TYPE OF APPLIED PART: |  |  |  |
| --- | --- | --- | --- | --- |
|  | What is the APPLIED PART for an MRI scanner? |  |  |  |
|  | a | The RF coil. |  |  |
|  | b | The gradient coil. |  |  |
|  | c | The table which supports the patient. |  |  |
|  | d | There isn’t one. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 10. | ### LEAKAGE CURRENT: |  |  |  |
| --- | --- | --- | --- | --- |
|  | The definition of LEAKAGE CURRENT includes the following: |  |  |  |
|  | a | Current that is functional |  |  |
|  | b | Current that is not functional |  |  |
|  | c | Current that flows when the EQUIPMENT is malfunctioning |  |  |
|  | d | PATIENT AUXILIARY CURRENT |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 11. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Leakage currents can arise due to the following: (Note this is probably not described in BS EN 60601-1:2006) |  |  |  |
|  | a | Inductors in mains-inlet filters |  |  |
|  | b | Capacitors in mains-inlet filters |  |  |
|  | c | The capacitance between the (mains connected) primary of a transformer and the metal chassis of the transformer |  |  |
|  | d | The capacitance between the live and earth wires of the mains cable |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 12. | ### EARTH LEAKAGE CURRENT: |  |  |  |
| --- | --- | --- | --- | --- |
|  | The definition of EARTH LEAKAGE CURRENT includes the following: |  |  |  |
|  | a | Current flowing from the PATIENT through or across the insulation into the PROTECTIVE EARTH CONDUCTOR |  |  |
|  | b | Current flowing from the APPLIED PART through or across the insulation into the PROTECTIVE EARTH CONDUCTOR |  |  |
|  | c | Current flowing from the ENCLOSURE through or across the insulation into the PROTECTIVE EARTH CONDUCTOR |  |  |
|  | d | Current flowing from the MAINS PART through or across the insulation into the PROTECTIVE EARTH CONDUCTOR |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 13. | ### PATIENT LEAKAGE CURRENT: |  |  |  |
| --- | --- | --- | --- | --- |
|  | The definition of PATIENT LEAKAGE CURRENT includes the following: |  |  |  |
|  | a | Current flowing from the ENCLOSURE through the PATIENT into the PROTECTIVE EARTH CONDUCTOR |  |  |
|  | b | Current flowing from the PATIENT CONNECTIONS via the PATIENT to earth. |  |  |
|  | c | Current originating from the unintended appearance of a voltage from an external source on the PATIENT and flowing from the PATIENT via the PATIENT CONNECTIONS of an F-TYPE APPLIED PART to earth. |  |  |
|  | d | Current flowing from the MAINS PART through or across the insulation into the PROTECTIVE EARTH CONDUCTOR |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 14. | ### PROTECTIVE EARTH TERMINAL: |  |  |  |
| --- | --- | --- | --- | --- |
|  | The PROTECTIVE EARTH TERMINAL is a terminal connected to conductive parts of: |  |  |  |
|  | a | Type BF APPLIED PARTS |  |  |
|  | b | TYPE CF APPLIED PARTS |  |  |
|  | c | CLASS I equipment for safety purposes. |  |  |
|  | d | CLASS II equipment for safety purposes. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 15. | ### CONTROLS |  |  |  |
| --- | --- | --- | --- | --- |
|  | A temperature sensing control that is intended to keep a temperature within a specific range or above/below a preset value is: |  |  |  |
|  | a | A thermostat |  |  |
|  | b | A thermal cut-out |  |  |
|  | c | An operator settable thermal cut-out |  |  |
|  | d | An over-current release |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 16. | ### SINGLE FAULT CONDITIONS |  |  |  |
| --- | --- | --- | --- | --- |
|  | In relation to protection against electrical hazards from ME EQUIPMENT, which of the following are defined as SINGLE FAULT CONDITIONS in BS EN 60601-1:2006 |  |  |  |
|  | a | Open circuit of a PROTECTIVE EARTH CONDUCTOR. |  |  |
|  | b | Open circuit of a FUNCTIONAL EARTH CONDUCTOR. |  |  |
|  | c | Earthing a PATIENT. |  |  |
|  | d | Short circuit of any one CREEPAGE DISTANCE or AIR CLEARANCE that complies with the requirements for one MEANS OF PROTECTION. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 17. | ### GENERAL TEST REQUIREMENT |  |  |  |
| --- | --- | --- | --- | --- |
|  | The tests described in BS EN 60601-1:2006: |  |  |  |
|  | a | Are type tests. |  |  |
|  | b | Are routine tests which should be carried out periodically on all ME EQUIPMENT. |  |  |
|  | c | Are potentially destructive. |  |  |
|  | d | Include tests which need not be performed if analysis shows that the condition being tested has been adequately evaluated by other tests or methods. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 18. | ### CLASSIFICATION |  |  |  |
| --- | --- | --- | --- | --- |
|  | ME EQUIPMENT and its APPLIED PARTS is classified by marking and/or identification as described BS EN 60601-1:2006. Which of the following does this include: |  |  |  |
|  | a | The CLASS of protection against electric shock |  |  |
|  | b | The degree of protection against electric shock |  |  |
|  | c | The methods of sterilization or disinfection recommended by the manufacturer |  |  |
|  | d | The mode of operation |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 19. | ### COLOURS OF THE INSULATION OF CONDUCTORS |  |  |  |
| --- | --- | --- | --- | --- |
|  | Any insulation on conductors inside ME EQUIPMENT that form PROTECTIVE EARTH CONNECTIONS shall be identified: |  |  |  |
|  | a | Throughout its length by the colours green and yellow. |  |  |
|  | b | By the colours green and yellow at least at the termination of the conductors. |  |  |
|  | c | Throughout its length by the colour green. |  |  |
|  | d | By the colours brown at least at the termination of the conductors. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 20. | ### INDICATOR LIGHTS |  |  |  |
| --- | --- | --- | --- | --- |
|  | For indicator lights the colour red can be used on ME EQUIPMENT |  |  |  |
|  | a | To indicate that the mains power in ON. |  |  |
|  | b | For any purpose defined according to the manufacturer. |  |  |
|  | c | Exclusively to indicate a warning of danger and immediate action required. |  |  |
|  | d | To indicate that the EQUIPMENT should not be used. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 21. | ### DISPLAYS |  |  |  |
| --- | --- | --- | --- | --- |
|  | Dot-matrix and other alphanumeric displays: |  |  |  |
|  | a | Are considered to be indicator lights |  |  |
|  | b | Are not considered to be indicator lights |  |  |
|  | c | Are considered to be indicator lights if they are lit continuously |  |  |
|  | d | Are only considered to be indicator lights if they are red and lit continuously |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 22. | ### DOCUMENTS |  |  |  |
| --- | --- | --- | --- | --- |
|  | Which of the following constitute a minimum set of documents which should accompany ME EQUIPMENT |  |  |  |
|  | a | Instructions for use. |  |  |
|  | b | A technical description |  |  |
|  | c | An address to which the user can refer |  |  |
|  | d | A circuit diagram for the EQUIPMENT |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 23. | ### POWER SUPPLY |  |  |  |
| --- | --- | --- | --- | --- |
|  | ME EQUIPMENT shall be suitable for a power supply having no voltage in excess of the nominal value plus: |  |  |  |
|  | a | 5% |  |  |
|  | b | 10% |  |  |
|  | c | 15% |  |  |
|  | d | 20% |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 24. | ### EXTERNAL D.C. POWER SOURCE |  |  |  |
| --- | --- | --- | --- | --- |
|  | In ME EQUIPMENT specified for power supply from an external d.c. source (e.g. an ambulance) if a connection with the wrong polarity is made: |  |  |  |
|  | a | The EQUIPMENT must still function correctly. |  |  |
|  | b | No SAFETY HAZARD, other than absence of ESSENTIAL PERFORMANCE shall develop. |  |  |
|  | c | The EQUIPMENT shall be rendered inoperable. |  |  |
|  | d | A warning lamp should display the reversed polarity condition. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 25. | ### PLUG |  |  |  |
| --- | --- | --- | --- | --- |
|  | ME EQUIPMENT intended to be connected to the SUPPLY MAINS by means of a plug shall be designed such that 1 second after disconnection the voltage between the supply pins does not exceed: |  |  |  |
|  | a | 20V |  |  |
|  | b | 40V |  |  |
|  | c | 60V |  |  |
|  | d | 80V |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 26. | ### TESTING (IN-)ACCESSIBILITY OF LIVE PARTS |  |  |  |
| --- | --- | --- | --- | --- |
|  | Accessibility of LIVE parts through an opening of an ENCLOSURE shall be tested as follows: |  |  |  |
|  | a | A straight unjointed test finger. |  |  |
|  | b | A test rod of diameter 4mm and length 100mm. |  |  |
|  | c | A “mains-tester” screwdriver. |  |  |
|  | d | For Category AP and APG ME EQUIPMENT there shall be no opening in the top of an ENCLOSURE. |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 27. | ### PROTECTIVE EARTHING |  |  |  |
| --- | --- | --- | --- | --- |
|  | For ME EQUIPMENT without a POWER SUPPLY CORD the impedance between the PROTECTIVE EARTH TERMINAL and any ACCESSIBLE METAL PART which is PROTECTIVELY EARTHED shall not exceed: |  |  |  |
|  | a | 0.1 |  |  |
|  | b | 0.15 |  |  |
|  | c | 0.2 |  |  |
|  | d | 1 |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 28. | ### PROTECTIVE EARTHING |  |  |  |
| --- | --- | --- | --- | --- |
|  | For ME EQUIPMENT with a non-detachable POWER SUPPLY CORD the impedance between the protective contact in the MAINS PLUG and any METAL PART which is PROTECTIVELY EARTHED shall not exceed: |  |  |  |
|  | a | 0.1 |  |  |
|  | b | 0.15 |  |  |
|  | c | 0.2 |  |  |
|  | d | 1 |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 29. | ### PROTECTIVE EARTHING |  |  |  |
| --- | --- | --- | --- | --- |
|  | The impedance of the PROTECTIVE EARTH circuit is checked by a 50/60Hz current with a no load voltage of 6V passed for 5 to 10 seconds. To comply with 60601-1:2006 the current should be: |  |  |  |
|  | a | 1A |  |  |
|  | b | 10A |  |  |
|  | c | 25A or 1.5 times the rated current of the EQUIPMENT, whichever is greater |  |  |
|  | d | Limited in value so as not to damage the EQUIPMENT |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 30. | ### LEAKAGE CURRENTS |  |  |  |
| --- | --- | --- | --- | --- |
|  | The specified maximum LEAKAGE current values for all types of leakage current apply under which of the following conditions: |  |  |  |
|  | a | In the NORMAL CONDITION |  |  |
|  | b | In the specified SINGLE FAULT CONDITIONS |  |  |
|  | c | With ME EQUIPMENT energised in stand-by condition and fully operating and with any switch in the MAINS PART in any position |  |  |
|  | d | With a supply equal to 110% of the highest RATED MAINS VOLTAGE |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 31. | ### PATIENT LEAKAGE CURRENT |  |  |  |
| --- | --- | --- | --- | --- |
|  | The PATIENT LEAKAGE CURRENT is measured in TYPE CF APPLIED PARTS: |  |  |  |
|  | a | From and to every PATIENT CONNECTION in turn |  |  |
|  | b | Between any single PATIENT CONNECTION and all other PATIENT CONNECTIONS connected together |  |  |
|  | c | From and to all PATIENT CONNECTIONS of a single function connected together |  |  |
|  | d | Through the PATIENT |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 32. | ### MEASURING DEVICE |  |  |  |
| --- | --- | --- | --- | --- |
|  | The measuring device used to measure LEAKAGE CURRENTS should have a resistive impedance of approximately: |  |  |  |
|  | a | 100 |  |  |
|  | b | 1k |  |  |
|  | c | 10k |  |  |
|  | d | 1M |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 33. | ### MEASURING DEVICE |  |  |  |
| --- | --- | --- | --- | --- |
|  | The measuring instrument used to measure LEAKAGE CURRENT which forms part of the measuring device should measure: |  |  |  |
|  | a | Resistance |  |  |
|  | b | RMS current |  |  |
|  | c | RMS voltage |  |  |
|  | d | Peak current |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 34. | ### INSULATION TEST |  |  |  |
| --- | --- | --- | --- | --- |
|  | Which insulation test does BS EN 60601-1:2006 specify? |  |  |  |
|  | a | Insulation resistance |  |  |
|  | b | Dielectric strength |  |  |
|  | c | Electrical resistance |  |  |
|  | d | Mechanical strength |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 35. | ### APPLIED PART |  |  |  |
| --- | --- | --- | --- | --- |
|  | On which of the 3 types of APPLIED PART is it appropriate to carry out the MAINS VOLTAGE on the APPLIED PART – PATIENT LEAKAGE CURRENT TEST? |  |  |  |
|  | a | None of them |  |  |
|  | b | TYPE B |  |  |
|  | c | TYPE BF |  |  |
|  | d | TYPE CF |  |  |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

|  | ### MAXIMUM ALLOWABLE LEAKAGE AND PATIENT AUXILIARY CURRENT |  |
| --- | --- | --- |
|  | According to BS EN 60601-1:2006, what are the maximum allowable continuous LEAKAGE and PATIENT AUXILIARY CURRENTS in the following Questions (35 to 50)? Select one of the answers indicated by the letters a to k below: |  |
|  | a | 1A |
|  | b | 10A |
|  | c | 50A |
|  | d | 100A |
|  | e | 500A |
|  | f | 1mA |
|  | g | 5mA |
|  | h | 10mA |
|  | i | None of the above |
|  | j | Two fault conditions are present and only NC and SFCs should be tested |
|  | k | Not specified in this “part” of the standard i.e. BS EN 60601-1:2006 |

| 36. | ### The EARTH LEAKAGE CURRENT of any general type of ME EQUIPMENT if one supply conductor is interrupted. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 37. | ### The EARTH LEAKAGE CURRENT of a Type CF Cardiac Monitor when normally connected. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 38. | ### The TOUCH CURRENT of a Type B Lung Function Monitor if the PROTECTIVE EARTH is interrupted. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 39. | ### The TOUCH CURRENT of a normally connected Type BF ECG Recorder. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 40. | ### The TOUCH CURRENT of a Type CF Cardiac Monitor which has reversed mains connections. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 41. | ### The a.c. PATIENT LEAKAGE CURRENT of a Type BF Cardiotachograph with a reversed mains supply. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 42. | ### The a.c. PATIENT LEAKAGE CURRENT of a Type B Ultrasonic Fetal monitor with an interrupted PROTECTIVE EARTH. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 43. | ### The a.c. PATIENT LEAKAGE CURRENT of a Type CF Monitor when normally connected. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 44. | ### The a.c. PATIENT LEAKAGE CURRENT of a Type CF Monitor when a supply conductor is interrupted. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 45. | ### The a.c. PATIENT LEAKAGE CURRENT of a Type BF Monitor when normally connected but with mains on one of the patient connections. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 46. | ### The a.c. PATIENT LEAKAGE CURRENT of Type CF APPLIED PARTS with the PROTECTIVE EARTH connector interrupted and mains on one of the patient connections. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 47. | ### The d.c. PATIENT AUXILIARY CURRENT for any Type of ME EQUIPMENT with the PROTECTIVE EARTH connector interrupted. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

| 48. | ### The a.c. PATIENT AUXILIARY CURRENT for a Type CF Cardiac Monitor which is normally connected. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer | Key words in BSEN 60601-1 index | Clause, table or figure | Comments |

The final two questions, 49 and 50, refer to figures a to e on the following page.

| 49. | ### The Instrument Under Test (IUT) is Class 1, with a Type BF Applied Part. Which, if any, of the configurations in figures a to e could be used to measure the “Earth Leakage Current”? |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer |  |  |  |

| 50. | ### The Instrument Under Test (IUT) is Class 1, with a Type B Applied Part. Which, if any, of the configurations in figures 1 to 5 could be used to measure the “Patient Leakage Current (Mains on Applied Part)”? |  |  |  |
| --- | --- | --- | --- | --- |
|  | Answer |  |  |  |

MP

AP

IUT

MD

ENCLOSURE

L

N

MP

AP

IUT

ENCLOSURE

L

N

MD

Figure a					Figure b

MP

AP

IUT

MD

ENCLOSURE

L

N

MP

AP

IUT

MD

ENCLOSURE

L

N

1:1.1

TRANSFORMER

MP

AP

IUT

MD

47k

ENCLOSURE

L

N

L

N

Figure c					Figure d

Figure e