Resonant Cavities

as Beam Position Monitors

Part 3. Analog signal processing.

Laboratory measurements.

A. Liapine

This part explains the basic principles used in the electronics for the BPM’s signals processing. Two basic methods of the downconversion, heterodyne and homodyne,  are discussed. Simple tests which can be done to the electronics and to the system BPM+electronics are described briefly.

- Electronics

Signals coming from a cavity beam position monitor usually oscillate at a very high frequency – from a few GHz up to a few tens of GHz. These signals must be downconverted before they may be digitized for the following analysis, because the capabilities of the digitizers available on the market are still limited to a few hundreds of MHz. Basically, two methods are used in the radio technics – heterodyne and homodyne. They also can be used for the conversion of the signal to a higher frequency, but we are not interested in this possibility and furthermore – need to suppress the high harmonic. We will start with the heterodyne receiver and then continue with the homodyne one (which could be treated as an extremely simplified version of the heterodyne unless it would not have some special properties).

  - Heterodyne Receiver

The (super-)heterodyne receiver is widely spread one. It is used everywhere starting from personal communication devices to radio and TV tuners. Figure below illustrates the main principle of heterodyne receiver. Radio frequency (RF) signal is first amplified in a frequency selective (but usually broadband) low noise stage, then translated to a lower intermediate frequency (IF): fIF = fRF – fLO. After a significant amplification and an additional filtering on the intermediate frequency it is finally downconverted to the baseband (i.e. to the original or desired signal frequency).

LO

Band select filter

Mirror frequency rejection filter

Channel select filter

RF

IF

f

Band select

filter

LO

LO +900

f

Channel select filter

fIF

"I"

"Q"

LNA

This technique is flexible in changing the receiving frequency, because the only change to be done is the frequency of the first local oscillator (LO) in the way that at the frequency of the signal at the output of the mixer would not change. The rest of electronics is free from additional readjustments, what is very important in the applications of the heterodyne receiver in the radio, TV, satellite and other communications.

The 2-stage downconversion of the BPM signal is sometimes needed to reduce the losses introduced by long cables carrying the BPM signals out of the accelerator tunnel. In that case the first stage of downconversion is placed to the BPM as close as possible, the signal is downconverted to some intermediate frequency, cable losses for which are considerably lower. The second stage is then placed outside of the tunnel, in the region, which is not as strongly loaded with radiation of every kind.

The main problem in heterodyne technique is so-called image frequency. To understand the problem let's first discuss in more detail the principle of the mixer’s functionality.

The RF-signal at the front end of the electronics oscillates with some frequency RF,

URF(t)=acos(ωRFt)							(1)

The signal coming from the local oscillator has some fixed frequency LO

ULO(t)=bcos(ωLOt)							(2)

These two signals are multiplied by a non-linear element of the mixer. As a result two signals with frequencies RF - LO  and RF + LO are produced.

URF(t)ULO(t)=ab2[cos(ωRF−ωLO)+cos(ωRF+ωLO)]	(3)

Taking the first term we have a down-conversion. The second term corresponds to up-conversion. Usually we are interested in the downconverted term and ωLO>ωIF. In this case a frequency of ωIM=ωLO−ωIF is downconverted to a negative IF frequency. In this way the signal at this frequency joins the signal of interest. Therefore any noise or interference signals at this frequency have to be rejected. The frequency ωIM is called image frequency.

To reject the mirror frequency signal an additional filter is often applied in front of the mixer.

ωLO

ωIM

ωRF

ωIF

ωIF

RF, IM

LO

IF

0

-ωIF

ωIF

f

f

1.2 Homodyne receiver

Nowadays in a close connection to the growing mobile and wireless communications more and more attention is paid to direct conversion receivers. Direct conversion reception, also referred to as homodyne, or zero-IF conversion, is the most natural solution to detect information transmitted by a carrier in just one conversion stage.

In this case the electronics has at least one mixer stage less then in a heterodyne system. But from the other hand the flexibility to change receiving frequency is reduced significantly, because the narrow band filtering is not possible before downconversion and the only narrow band filter is the output lowpass. Though, in case of the BPM signal processing there is no need to change the input frequency. That means a narrow band filter can be used already at the front-end of the electronics.

The principle is illustrated below. The signal is first amplified at a low noise stage and then directly converted to the baseband or even to a direct current signal. When the frequencies of the RF and the LO signals are equal, this scheme works as a phase detector. In some literature, only when the local oscillator is synchronized in phase with the incoming carrier frequency, the receiver is called homodyne.

Band select

filter

DC

f

Band select

filter

LO

LO +900

f

Low-pass

Filter

“I”

“Q”

Low-pass

filter

LNA

Low-pass

filter

Suppose that the IF in a heterodyne is reduced to zero. The LO will then translate the centre of the desired channel to 0Hz, and the portion of the channel translated to the negative frequency half-axis becomes the “image” to the other half of the same channel at the positive frequency half-axis.

f

0

To achieve maximum information, we should take both parts of signal. It’s done by a method, which is called quadrature downconversion. The principle of this method is that the signal is at first divided into two channels and then downconverted by an LO signal, which has a phase shift of 90o in one channel with respect to another. The vector of the resulting signal is described as:

Signal=I2+Q2		arg(Signal)=ϕ=arctgQI

The main problem in homodyne technique is an offset caused by the LO signal leakage to the RF port of the mixer. The propagated signal reflects from the components in the front-end of the receiver and goes back to the mixer, where it is mixed down to DC. The offset can be considerable with respect to the signals to be measured. This leads to a narrower dynamic range of the electronics, because the active components get saturated easier than it would be in case of a zero offset.

For example, let’s take a mixer with the LO-drive equal to +10dBm and RF/LO isolation equal to 40 dB. In this case the offset can be as high as –30dBm, or about 2mV. In case of a high sensitivity this can be a large number (this is the signal level at the output of the mixer, some amplification stages follow!).

Band select

filter

DC

Low-pass

filter

LO

LNA

Another problem of the homodyne receiver, or, more concretely, of the I/Q (in-phase/quadrature) mixer, is mismatches in its branches. Assuming a mismatch of ε for the amplitude and θ for the phase, we can estimate the error, caused by these mismatches. In this way we get:

EIQ=Sideal−SmissSideal≈12+εθ2.

For typical values ofε=0.3 and θ=30 this gives an error of 1.5⋅10−3.

“I”

“Q”

In the processing of the BPM signals it is important to know the phase with respect to the phase of the reference cavity. That means that an I/Q mixer becomes very attractive for this application. On the other side, it has too many specific hardware problems. That’s why it is desirable to apply a digital downconversion at the last stage of the downconversion. That means that a processing usually applied to the signal in an analog form, is applied to a digitized signal in a digital form. In that case an I/Q mixer and other elements realized in forms of idealized models have no mismatches or reflections etc. Surely, the initial signal still has to be downconverted and prepared for the digitizing. The entire scheme looks then like a heterodyne scheme in the picture at the very beginning of this part with the difference that the second downconversion stage is not real but realized in digits.

- Electronics Tests

As the entire electronics relays on mixers, it makes sense to test the mixers before they are inserted into the electronics. The most important properties of the mixers are the transmission characteristic (linearity, 1dB compression point, conversion loss), isolation and, perhaps, higher harmonics.

The transmission characteristic can be measured with two oscillators, attached to RF and LO inputs of the mixer and any device measuring the signal level – spectrum analyzer, oscilloscope, ADC, even simple voltmeter. In case of oscilloscope the form of the signal can be controlled visually for any distortions, while a spectrum analyzer shows distortions due to non-linearity directly as a set of harmonics.

G1

G2

SA

RF

LO

IF

The signal level of the oscillator attached to the LO port of the mixer is kept constant and equal to the nominal level required for the mixer’s operation. The signal level at the RF port is changed and the level of the output signal is controlled. This test allows selecting the best mixers available for use in the electronics – ones with the best linearity, lowest losses, and highest input signal limit. In a very similar way the entire electronics can be tested as well.

The isolation between RF and LO ports can be measured directly by measuring the transmission from the LO to the RF port with a network analyzer. Isolation between LO and RF is very important for the mixers working in low frequency stages.

NA

50Ω

LO

RF

IF

G1

G2

X

Referenz

"I"

1,5GHz-

Elektronik

"Q"

"Ladung"

Oszilloskop