|  | SCT Fibre Optic Cable PRR2 |  |  |
| --- | --- | --- | --- |
| ATLAS Project Document No: | Institute Document No. | Created : 03/11/2003 | Page: 1 of 5 |
| ATC-TB-ER-0030 |  | Modified : 19/10/2004 | Rev. No .: 3.0 |

| Cable Production Readiness Review SCT FIBRE OPTIC CABLES Abstract The SCT readout uses 12 way ribbon fibres. The SCT FSI alignment system uses 8 way ribbon fibres. The fibres will be cabalised into a protective cable to ensure that they can be installed without damage. This updated version includes the description of the protection to be used for the bare ribbons at the ends of the cables. |  |  |
| --- | --- | --- |
| Prepared by: P. Wells (pippa.wells@cern.ch) C. Menot M. Stodulski | Checked by: T. Weidberg N. Unno S. Malyukov | Approved by: John Hill, Paul Coe, Alexander Grillo, Marco Olcese, Mike Tyndel, Steinar Stapnes, Kathy Pommes, Philippe Farthouat Friedrich Szoncso Gianpaolo Benincasa |
| Distribution List Marzio Nessi, Balazs Szeless |  |  |

# Introduction

The technical specifications, connectors and preliminary cable lengths of fibre-cables for the ATLAS SCT are described in the PRR1 document (EDMS,  ref. ATC-TB-ER-0015) which was issued prior to launching the market survey. The present document constitues the Production Readiness Review, and specifies the cable performance, procurement  procedure and quality control plan, cable length, labelling and installation issues.

# Cable function

Two types of cable are concerned:

- SCT Readout cable (6 and 8 times 12-way ribbons cabalised together)

- SCT FSI alignment cable (6 and 7 times 8-way ribbons cabalised together)

The SCT uses fibre ribbons for the readout of the data from the on-detetcor modules to the RODs as well as for the transmission of the TTC data from the RODs to the on-detector modules. The SCT FSI alignment system also uses fibre ribbons to carry laser light from USA15 to the SCT and vice versa. The ribbons are assembled into protective cables (‘cabalised’) to ensure that they can be installed without damage.

# Cable specifications

## Cable manufacturer

Fujikura Co. Ltd.

1-5-1 Kiba, Kohtoh-ku

Tokyo 135-8512

Japan

Contact: Y. Kurosawa

Telcommunication division

Telephone: +81-3-5606-1204, FAX: +81-3-5606-1536

e-mail: kurosawa@fujikura.co.jp

## Manufacturer data sheet

TECHNICAL DATA FOR TRIAL CABLE INCLUDING RADIATION TOLERANT FIBERS: see file B-03K0010.pdf included in this meta-document.

## User Specifications

The read out cables will be terminated by Fujikura with MT-8 and MT-12 ferrules supplied by NTT Advanced Technology Co. Ltd., AT-8MT-FER-GI and AT-12MT-FER-GI. The MT connectors require very specialised grinding tools and skilled operators. It is widely recognised that this should always be done in the factory and not in the field.

Most read out cables will have MT-12 connectors at both ends. These are denoted RO12/12.

One third of the barrel readout cables will have an MT-12 connector at one end (at the ROD crate in USA15), and two MT-8 connectors at the other end (the end closest to the detector). The 12-fibre ribbon will be spilt into two 6-fibre ribbons near one end. Each 6-fibre ribbon should then occupy the middle 6 holes of an MT-8 connector. This is sketched in Figure 1. These cables are denoted RO12/8+8. The MT-8 connectors in the split ends are interchangeable with MT-12 connectors as long as the 6 fibres are set in the 6 centre slots. The use of MT-12 connectors replacing the MT-8 ones is an option for Fujikura.

An Infineon spring should be placed over one end of each ribbon before the MT-12 connector is added. The spring should be located near one end of the straight MT-12 to MT-12 ribbons and near the MT-12 connector for the ribbons that split to two MT-8 connectors. The spring are standard "magnetic" types as the MT-12 ends of fibres are to be placed in USA15.

The FSI alignment cables will not be terminated with connectors by Fujikura. Instead, they will be fusion spliced to fibres coming from the detector and the racks in the ATLAS pit and to the racks in USA15.

Figure 1: Sketch of the connector and spring arrangement for cables denoted RO12/8+8

## Safety Requirements

Fire safety tests have been carried out by the supplier, as listed on page 3 of the document referred to in section 3.2.

CERN TIS confirms that the fibre cable does not contain halogen or sulphur, and passes the IEC 60332-1 test ( flame propagation). The cable conforms to the IS 23. Ref. TIS TEST REPORT : TIS-2003-64.

# Cable lengths, labels and installation

## Cable lengths

Cable lengths are provided in the following tables for the two sides of the ATLAS detector (denoted side A and side C). These tables, and accompanying tables with more complete details of the length calculation, are included in the file SCT_fibre_length_10Feb2004.xls attached to this meta document. Fujikura are not required to cut the alignment cables to length, but just to manufacture sufficient cable length to satisfy these requirements. The alignment fibre should therefore be supplied as

- 1 drum of 7-ribbon cable, up to 800m but not less than 670 m

- 1 drum of 6-ribbon cable, up to 2660m but not less than 2480 m

Worksheet “Side A summary”  from attached excel file

Table 1: Numbers of cables of each type on Side A with their lengths, ATLAS labels and ATLAS unique numbers

Worksheet “Side C summary”  from attached excel file

Table 2: Numbers of cables of each type on Side C, with their lengths, ATLAS labels and ATLAS unique numbers

## Spare cables

Spare cables should be provided as follows, subject to availability of cabalised-ribbon:

- Spare alignment cable should not be cut to specific lengths, since these cables do not have connectors attached and can be cut as needed. The 7-ribbon cables will be installed one per quadrant, and will thus supply one spare ribbon per quadrant. These cables have the same outer diameter as the 6-ribbon cables, so this has no impact on the layout.

- Two spare readout cables of the longest length requested should be made, with connectors, one of type RO12/12-8s, and one of type RO12/8+8-8s.

It should be noted that 8s RO cables can be used to replace 6s RO cables if necessary. Spare ribbons with connectors are being made for the PPF0 to PPF1 region. These spares may equally well be used to replace any damaged ends of the RO cables by splicing.

## Cable and ribbon labels

Tables 1 and 2 also include the ATLAS labelling convention. Labels should be attached to the cables and ribbons as follows:

The manufacturer should print a label defining the cable type along the outer casing, every 5 m:

- Definition of the cable type, eg: O(S)/10.5-8s-12/RH, O(S)/10.5-6s-8/RH

where O(S) stands for Optical fibre-cable for SCT, 10.5 for the outer diameter, 8s or 6s for the number of ribbons, 12 or 8 for the number of fibres in a ribbon, and RH for radiation-hard.

The two drums of alignment fibre-cable should also be clearly labelled on the outside of the drum.

For readout cables that are cut to length by Fujikura, the manufacturer should put labels on each end of each of  the cables around the outer casing. These labels should include three items:

- Definition of the cable start and end points, eg. ISE/A.F2.L2/RoT/Y.22-11.A2.

- Bar code of the ATLAS unique number of the cable, eg. 2005001, in code128 format

- A human-readable print of the ATLAS unique number, eg. 2005001.

The distances from the cable ends to the labels are also indicated in the tables. These distances DO NOT IMPLY that the outer casing should be stripped off all the way to the labels. The amount of outer casing to be stripped off will be decided during installation in the ATLAS pit.

The distances at the detector end (MT-8 end of cables with split ribbons) are

- 5m for barrel readout cables

- 2.5m for endcap readout cables

- 1.5m for alignment cables (to be cut to length eventually at CERN)

At the rack end, (the end with the spring for readout cables) the distance from the cable end to the label is

- 1.5m for all cables.

The manufacturer should put labels defining the ribbons on each end of the ribbons. These labels should show:

- For RO12/12 ribbons, barcode of ATLAS unique number of the cable plus 01 through 08 (or 06) for the ribbon number, eg. 200500101, 200500102… in code128 format, together with a human-readable print of 200500101…

- For RO12/8+8 split ribbons, the MT-12 ends should be labelled with the cable number plus 10, 20,…, e.g., 200508510, and the MT-8 ends with the cable number plus 11 and 12, 21 and 22, …, e.g., 200508511 and 200508512, to indicate the two halves of each of the 6 or 8 ribbons.

## Cable Installation

The data cables will be supplied with connectors at both ends. Surplus lengths of each cable will be stored in USA15 in the false floor under the racks. The FSI cables will be fusion spliced at both ends.

## Cable End Protection

The data fibre cables will be supplied with conenctors at both ends and the ends will therefore need to be protected during cable pulling. It is anticpated that the cable will be pulled from the top of the tile calorimeter towards USA15. Therefore the end that is being pulled towards USA15 will require stronger protection. The “light duty rigging” end protection at the end of a cable is shown in the photograhs in Figure 2, Figure 3 and Figure 4. The pull strength of this “light duty rigging” is greater than 220 N which is greater than the maximum allowed pull strength for the cable itself. The outer diameter of the protection tube is 35 mm for a cable with “straight” ribbons and 50 mm for a cable with “Y junction” ribbons.  The ends of the cables which are being laid in from the top of the calorimeter to the PPB1 and PPF1 patch panels do not require this type of protection. Instead simple ”air bubble sheet” protection will be used. Therefore the “light duty rigging” should be used on the end of the cables which have the Infineon springs over the ribbons and the “air bubble sheet” should be used for the ends which do not have the Infineon springs.

Figure 2Fibre cable end protection

Figure 3 Fibre Cable end protection showing transition to 10.5 mm diameter.

Figure 4 Fibre cable end protection. View of end showing chain for pulling.

# Procurement procedure and quality assurance plan

## Quality control tests

The attenuation of the ribbon fibre for the readout is less than 20 dB/km. The connector insertion loss is less than 0.5 dB/connector. The attenuation loss of the fibres in the cables after MT-connector termination in both ends is less than 1.0 + 20 * L[km] dB. All fibres of the readout cables shall be tested by the manufacturer and to pass the afore-mentioned value.

The alignment fibre ribbons should be tested for attenuation at 800 nm and 850 nm after cabalisation, as they were after manufacture, to ensure that no alignment fibre cable has attenuation greater than 15 dB per km.

## Pre-production tests

A sample of 5 m trial cable has been submitted by the manufacturer and verified at CERN.

## Quality assurance plan

The following items of the delivered readout cables shall be measured and documented by the manufacturer:

1) Number of cables

2) Labelling of the ribbons and cables

3) Lengths of the ribbons in the cables

4) Attenuation loss of all fibres in the cables from connector to connector

The following items of the delivered alignment cables shall be measured and documented by the manufacturer:

1) Number of cables

2) Labelling of the ribbons and cables

3) Lengths of the ribbons in the cables

4) Attenuation loss of all fibres in the cables from end to end