## Slide 1: Fixed Frequency Trigger Veto
The problem:
  Currents in wire bonds in presence of strong magnetic fields
  DC current not a problem (small force)
  Time-varying current a problem if frequency matches resonance condition for wire-bond
  Observed in CDF SVXII detector (lost modules)
  Main concern for SCT is bonds carrying current from VDC to VCSELs. eg data after a trigger
    Equivalent to 5 mA burst of current when triggers sent

## Slide 2: Resonance condition
Finite element analysis and physical tests show narrow resonance peak near 18 kHz (55 us) for 3mm long bond wires with plane of wire bond perpendicular to Lorenz force (the unfavourable direction).
The actual wire bonds in the SCT will have lower amplitude and higher frequency resonances.
For detector safety the SCT will avoid running with fixed frequency triggers with frequencies above 15 kHz (but below 500kHz)
N.B. resonance is narrow so effect on physics running should be small (see later)
FEA
Wire test
doi:10.1016/j.nima.2004.08.120
cdf-video

## Slide 3: Trigger sources
TIM
ROD
ROD
ROD
Modules
Local Trigger Processor
In calibration mode triggers can/will come from other sources
  LTP
  TIM
  ROD
Central Trigger Processor
SCT DAQ crate
L1A
BUSY
In physics mode triggers come from the CTP
Trigger veto logic runs back up the chain
External source
External source

## Slide 4: ROD mode
Used in most calibration scans
  Response curve, Noise occupancy, …
DSP code Master DSP controls trigger rate
Current implementation limits rate to < 15 kHz.
  Good for detector safety
  Does slow down calibration
Could allow faster triggers if pseudo-random
  Not simple, but possible
ROD
ROD
ROD
Modules

## Slide 5: TIM mode
TIM implements a Fixed Frequency Trigger Veto (FFTV) algorithm
Repetitive period monitoring
  Count clocks between triggers
  Store results
  Check for period match
  If match veto
Currently has several changeable parameters
  Possible to modify e.g. matching criteria, veto time, over VME
TIM
ROD
ROD
ROD
Modules
Ext

## Slide 6: Veto effect in TIM mode
TIM currently vetos fixed frequency triggers from either its internal generator or lemo sources when in calibration mode
  Randomiser is used to produce triggers during TIM calibration scans
    Tested with modules etc in SR1
  Application of veto generates warning for user

## Slide 7: Veto effect in LTP or CTP (physics) mode
Same algorithm detects FFT
TIM asserts BUSY to LTP
LTP (or CTP) should stop sending triggers
  What to do if it doesn't?
TIM
ROD
ROD
ROD
Modules
Local Trigger Processor
Central Trigger Processor
BUSY

## Slide 8: Remaining risks 1/3
TIM observes FFT but BUSY logic doesn't stop LTP/CTP from sending triggers
  Solution: after a short (2μs) delay for busy signal propagation, TIM will go into “emergency” mode. Refuses to send triggers.
  Clearance from this state requires manual intervention

## Slide 9: Remaining risks 2/3
Enthusiastic user disables trigger veto algorithm by changing registers and/or using jumper
  Solution: parameters will be hard-wired into firmware. No way to change without reflashing TIM.
  CAVEAT: Pixels! Will detect hardware ID of board, and if not an SCT board, will allow a jumper AND a register bit to disable the Veto.

## Slide 10: Remaining risks 3/3
Amateur ‘expert’ decides to send fixed frequency triggers directly from crate controller to ROD (not via MDSP histogramming code)
  No way for ‘user’ to do this (at present)
  More water-tight restriction required?
  (The FFTV algorithm is quite small might be able to be squeezed into the ROD local trigger firmware)

## Slide 11: Bullet-proof TIM firmware
Will include solutions to remaining issues (1/3) and (2/3) above
  Emergency mode
  Hardwired parameters
  Veto logic placed ‘last thing’ before triggers leave TIM
To be ready by early March?
Place holder for new logic diagram

## Slide 12: Effect on Physics Running
Provided ATLAS L1A’s are not sent at fixed frequencies in the ‘danger zone’ then veto application will be very rare (less than about 1/day)
If triggers are send at fixed frequencies then we really do want to apply the veto more often!
Place holder for veto rate plot