## Slide 1
ATLAS SCT/Pixel TIM FDR/PRR
28 June 2004
Physics & Astronomy HEP Electronics
Martin Postranecky John Lane, Matthew Warren
TIM HARDWARE

## Slide 2: COMPONENTS USED
2x FPGAs used :
FPGA-1 : Xilinx 2S200E-6FG456C
FPGA-2 : Xilinx 2S600E-7FG456C
2x Xilinx PROMs
7x Delay Lines PDU-16
4x Clock drivers/mpxs
15x TTL/PECL/ECL/NIM
14x LV buffers
38x various TTL
2x DC/DC modules
8x transistors
38x LEDs
626x passive comp.
20x connectors
80x pin headers

## Slide 3: TIM-3A MODULE
9U x 400 mm, single width, VME64x module
Standard VME slave interface with A24 / D16
or A32 / D16 access
A16-23 ( or A16-31) Base Address, or GA address
TIM-3A Power :
~ 7A of +5V ~ 2A of +3V3 ( incl. +1V8 supply)
< 1A of +12V
} for -5V2 supply
< 1A of -12V
Total power ~ 120 W

## Slide 4: TIM Discretes vs FPGAs
Notes: 2x FPGAs used :
FPGA-1 : Xilinx 2S200E-6FG456C
FPGA-2 : Xilinx 2S600E-7FG456C
2x Xilinx PROMs
7x Delay Lines PDU-16
4x Clock drivers/mpxs
15x TTL/PECL/ECL/NIM
14x LV buffers
38x various TTL
2x DC/DC modules
8x transistors
38x LEDs
626x passive comp.
20x connectors
80x pin headers

## Slide 5: TIM Functional Model
Notes: TIM FUNCTIONAL MODEL shows the main blocks of TIM :
TTC INTERFACE with its TTCrx ASIC, producing all the
timing, trigger and command signals received from the FLT
while in the RUN MODE
STAND-ALONE generator of the same signals in the SA MODE.
STAND-ALONE SEQUENCER which allows long command
sequences to be down-loaded from the host in the SA MODE
THE EVENT FIFOS AND SERIALISERS to allow IDs of
multiple triggers to be stored and output sequentially
BACKPLANE MAPPING logic

## Slide 6: ATLAS-SCT “TIM” Schematics
MP/(MRMW) v1.0 17-06-04
ECLEXTCLK
NIMEXTCLK
EN
6x ECLEXTCOMMANDS
6x NIMEXTCOMMANDS
6
6x EN
6
6x
4x EN
4
Internal Stand-Alone Commands
4x
VME Commands
6
6
2
ECLEXTBUSY
NIMEXTBUSY
6
6x
Sync
CLOCK
6x
6
6x ECLDATAOUT
4x NIMDATAOUT
4
6
ECLCLKOUT
6x DATAOUT
SA BUSY
NIMCLKOUT
6
MPX-PLL
DELAY
TTC Interface
TTCrx
6x SACOMMANDS
Sequencer
8x32K RAM
0-7
Sink
8x32K RAM
0-7
VME ADDRESS BUS
VME READ/WRITE
VME READ
6x EN
6
6
2x EN
2
2
6x
2x
6x SEQCOMMANDS
2x SEQID
6x EN
6
6x
CLOCK
6x TTCCOMMANDS
6
CLK40DES1
2
2x EN
VME READ/WRITE
TTCFIBRE
2x FIFO
2x SERIALISER
2
ID
COUNTERS
2
CLOCK
MPX-PLL
CLOCKOUT
2x ID
CLOCK
8x F/F
6x TTCCOMMANDS
6x SACOMMANDS
2x ID
6x SEQCOMMANDS
2x SEQID
CLOCKOUT
TTCA(0-7)
8x F/F
TTCB(0-7)
DOUBLE 8x OUTPUTS TO P3
CLOCK
16x DIFFERENTIAL CLOCK OUTPUTS TO P3
6x DATAOUT
6x F/F
6
2
8x SINKDATA
80.1573MHz XTAL Osc.
4
2x
DELAY
8x8 MAPPING
RODBUSYOUT
16x RODBUSY INPUTS FROM P3
16
16x RODBUSYMASK
16x
BUSY MODULE
Notes: TIM SCHEMATICS
This slide shows the schematic representation of the circuitry of the STAND-ALONE generator, together with the suite of NIM and ECL inputs and outputs.
There is an 80.1573MHz X-tal oscillator producing an internal 40.0786MHz clock.
Either this, or any externally generated 40MHz clock can be selected to provide the STAND-ALONE clock.
The fast commands ( trigger and resets ) can be generated internally, either repetitively at a single selected frequency ( 50Hz - 600kHz ), with trigger also available as a random trigger with selected average frequency, or by VME commands from the host by writing to the COMMAND register.
All the TTC-type commands, plus the 40MHz clock, can also be input from external sources in either NIM or differential ECL standard, using front panel connectors.
All the STAND-ALONE commands are synchronised to the selected clock and send to the BACKPLANE MAPPING logic to produce the TTC(0-7) output signals.
In addition, the trigger command can be delayed by a fully programmable delay of up to 143 clock periods to accommodate the FE data pipeline length.
There is a further capability of providing a "burst" of a fully-programmable ( between 0-16k ) number of either external or internal triggers or calibration triggers.
The selection of any/all of the above externally input or internally generated signals is controlled by an VME accessed ENABLES register, allowing for a full flexibility.
The L1ID counter and register allows any FLT number to be down-loaded from the host interface in STAND-ALONE mode prior to issuing triggers.
The FLTN counter will increment with each subsequent trigger, loading the new number into the FLTN register. In the RUN mode, this register is loaded from the TTC.
Similarly, the BCID counter and register provides the suitably offset BCN for every trigger, in both STAND-ALONE and RUN modes.
The BCNTIN register stores the BCN received from the TTC during run.
Finally, the TTID is stored in the TTYPE register.
The SEQUENCER RAM can be loaded from the host with up to 32k long sequence of commands to allow for various tests of the system.
There is an matching SINK RAM which allows us to store up to 32k of commands being issued by the TIM to RODs. Again, this is invaluable in testing the hardware and timing.
The slide also shows the BACKPLANE MAPPING, together with the backplane outputs synchronising drivers, the backplane PECL clock drivers, the two FIFOs and serialisers for trigger L1ID, BCID and TTID data, plus the TTCrx receiver and interface. As you can see, three sets of signals from three sources - STAND-ALONE, TTC and SEQUENCER - are received by the MAPPING module.
Finally, the basic ROD BUSY system receives ROD BUSY signals from all RODs in the crate and provides a masked-OR output to the FLT in either NIM or TTL standard ( selectable by a link ).

## Slide 7: A) In the RUN MODE :
1) INPUTS :
From TTCrx interface module : CLOCK40DES1, TTC Fast Commands, ID Numbers
From backplane : 16x ROD BUSY signals are received from individual RODs by PHILIPS N74LVT16244B LV Bus receivers. The active-low ROD BUSY lines are pulled up to +3V3 on the TIM, thus indicating empty ROD slots as NOT BUSY.
2) OUTPUTS :
The 40MHz clock is distributed to all BOC cards as 16 individual differential PECL pairs via the J3 backplane. Driver chip is MOTOROLA MC100E111JC 1:9 PECL DIFFERENTIAL CLOCK DRIVER, which guarantee channel-to-channel skew to be below 50 ps.
All on-board differential tracks, and all the P3 backplane tracks, are designed as point-to-point balanced 100R tracks of the same length ( ie. all the backplane slots receive clock with no time skew )
Each line of the differential pair has a 270R load resistor at the transmitter, and is terminated by 100R between the differential pair at the BOC receiver end.
Eight active-low TTC(0-7) outputs are bussed to RODs on two separate backplane buses, each for
8x RODs. The devices chosen, PHILIPS N74ABT574D, are Advanced BiCMOS Bus Interface drivers, which guarantees VOL of below 0.55V while sinking 64mA, and VOH of above 2V while sourcing 32mA.
With TIM is slot 13 in the middle of the crate, the TTC(0-7)A bus is terminated on the backplane in slot 21, and the TTC(0-7)B bus in slot 5.

## Slide 8: B) In the STAND-ALONE MODE :
1) INPUTS :
From INTERNAl or EXTERNAL sources :
All Clock and Command Signals, including Busy, can be input as either NIM or differential ECL signals on the Front Panel
Same ECL external commands can also be input on the second half of this IDC dual connector. This allows one TIM to operate as a master & drive another TIM slave. At least four TIM slaves can be driven by one TIM master using a daisy-chain ribbon cable and removing the ECL differential termination on the first three modules
2) OUTPUTS :
a) On backplane, same as in RUN MODE
b) On Front Panel, additional interface outputs :
The complete set of six TTC(0-5) type commands, together with BUSY and CLOCK, is output on one half of dual 16-pin IDC connector as 8x differential ECL pairs.
C) SEQUENCER :
Additionally, all 8x TTC-type commands can be input from the SEQUENCER RAM

## Slide 9
TIM Backplane Interfaces
Notes: BACKPLANE INTERFACES :
This slide shows the standards adopted for the backplane interface.
The 40MHz clock is distributed to all BOC cards as 16 individual differential PECL pairs via the J3 backplane. The devices chosen for this task are the MOTOROLA MC100E111JC 1:9 PECL DIFFERENTIAL CLOCK DRIVERs, which guarantee channel-to-channel skew to be below 50 ps.
Each line of the differential pair has a 270R load resistor at the transmitter, and is terminated by 100R between the differential pair at the BOC receiver end.
The eight active-low TTC(0-7) outputs are bussed to RODs on two separate backplane buses, each for 8x RODs. The devices chosen, PHILIPS N74ABT574D, are Advanced BiCMOS Bus Interface drivers, which guarantees VOL of below 0.55V while sinking 64mA, and VOH of above 2V while sourcing 32mA.
With TIM is slot 13 in the middle of the crate, the TTC(0-7)A bus is terminated on the backplane in slot 21, and the TTC(0-7)B bus in slot 5.
Finally, the 16x ROD BUSY signals are received from individual RODs by PHILIPS N74LVT16244B LV Bus receivers. The active-low ROD BUSY lines are pulled up to +3V3 on the TIM, thus indicating empty ROD slots as NOT BUSY

## Slide 10: TIM Front-Panel LEDs
MRMW v1.0 15-05-04
Overall Reset
VME Error 1
-5V, -12V Power On
NOTES:
VE Shows a VME bus error (flash) OR geog-addr error (i.e. wrong slot).
VA Flashes when TIM is accessed (addressed) by VME correctly.
RB In Stand-Alone Mode TIM is always busy.
All LEDs (apart from power supplies) have a 60ms pulse stretcher for better visibility.
Stand-Alone Mode
Stand-Alone Clock On
TIM Busy
ROD Busy’s (1 per slot)
+5V, +3.3V, +12V Power On
VME Access 2
TTC Clock On
TTC Mode (Run Mode)
Overall ROD Busy 3
L1A signal issued
-5
-12
OR
VE
SA
SC
TB
CA
BR
SP
+5
+3
+12
VA
TT
TC
RB
LA
ER
FR
14
15
16
17
18
19
20
21
12
11
10
9
8
7
6
5
ROD BUSY
ECR signal issued
FER signal issued
CAL signal issued
BCR signal issued
Spare signal issued

## Slide 11
TIM Front Panel Interfaces
Notes: FRONT PANEL INTERFACES :
This slide show the standards adopted for the inputs and outputs of external clocks and commands. The NIM inputs and outputs are provided with LEMO 00 connectors, and ECL CLOCKS are input and output on two-pin LEMO 0B connectors.
The complete set of six TTC(0-5) type commands, together with BUSY and CLOCK, is output on one half of dual 16-pin IDC connector as 8x differential ECL pairs.
Same external commands can also be input on the second half of this IDC dual connector. This allows one TIM to operate as a master & drive another TIM slave. At least four TIM slaves can be driven by one TIM master using a daisy-chain ribbon cable and removing the ECL differential termination on the first three modules.
Few words about the reasons for adopting the ECL and NIM standards for external inputs and outputs, requiring the additional -5V2 supply. As you may know, the CERN VMEbus Steering Committee recommended, in their paper "RECOMMENDED PRACTICES FOR THE USE OF VMEBUS IN PHYSICS APPLICATION", that differential ECL is the preferred method for front panel connectivity, and NIM for single-ended coaxial connections. It is also the case, that, for historical reasons, most of the off-the-shelf equipment available in CERN for eg. System tests uses NIM and differential ECL. For those reasons, and to make TIM as widely usable as possible, we have decided to retain the ECL and NIM standards as currently used by all the CLOAC modules.
This required the provision of -5V2 supply. We have provided an on-board DC-DC converter providing -5V from +/-12V supplies. The DC-DC converter can be disabled by a single link. Thus the total -5V2 supplied circuitry, which is only used for tests with externally generated clock and commands, can be disabled

## Slide 12: TIM3 Clock Flow
ECLEXTCLK
ECLEXTCLK2
CLK0
CLK00
CLOCK40DES1
BCCLK1B
9x CLOCK40
8x CLOCK40
SW8
NIMCLKOUT
ECLCLKOUT1
ECLCLKOUT2
8x
F/F
TTCout(0-7)
U47 DL4
SW7
TIM Setup
TTC(7-0)A
TTC(7-0)B
ROD Setup
PECL Drivers
SW9
TRIGCLK
Trigger Window
INT_CLK
EXTCLKB
CLKIN1
CLKIN2
MCLK1
SACLKB
TTCCLK1B
ENSACLK
ENINTCLK
CLKINB1
U45
SACLKLED
U56
BCCLKLED
U39
CT(5:0)
PCLKB
U42
CLKINB2
DL2OUT
8x
F/F
DL2OUTB
U44
CLKINB4
DL4OUT
U50
DL4OUTB
DL1OUT
NIMEXTCLK
U33
U42
80Mhz Osc.
U412
U44
U44
U46 DL2
WD(5:0)
Setup
Delay
U69
DL
WS(5:0)
Size
U62
DL
SW10
Size Comp.
U61
DL
U63
DL
U44
U52
TIMCLK1L
TIMCLK2L
TIMCLK3L
TTCCLK2B
U52
TTCCLK2L
FPGA2
FPGA2
FPGA2
FPGA2
FPGA2
FPGA2
FPGA2
FPGA1
U42
MRMW/MP v2.0 11-05-04
U36
EXTCLKLED
EXTCLK
U402
U412
U38
CLK
MUX
1
U57
CLOCK40DES2
CLOCK40
BCCLK1B
U48 DL1
U58
CLK
MUX
2
TTCrm/rq
U51
Notes: TIM-3 CLOCKS FLOW :
This slide show the clock system of the TIM. In SA MODE, an 80.1573 MHz on-board crystal clock generator provides the required 40.0786 MHz INTERNAL CLOCK.
There is a provision for accepting 40MHz EXTERNAL CLOCK in either NIM or differential ECL standard. Required clocks can be selected in the ENABLES register.
When in the RUN MODE, the TTC supplies the machine clock via the TTCrx ASIC.
There is a couple of delay lines on TIM, allowing for independent adjustments of ROD SETUP ( for TTC outputs timing ) and TIM SETUP ( of internal timings ). These delay lines are set by DIL switches as 6-bits in 0.5 nsec steps, giving more than adequate coverage over the 25nsec clock.
There is an additional delay adjustments available for the STAND-ALONE clock, set by the CLOCK DELAY register as 6-bits of 0.5nsec steps, which can be used to scan.
There are four additional delay lines, used to provide a TRIGGER WINDOW width and delay, which can be used with to reject random triggers

## Slide 13: TIM-3 Clock Switching/TIM-OK
MRMW/MP v2.0 18-06-04
INTCLK
TIMCLK
ENEXTCLK
ENTTCCLK
RUNMODE
TIM-OK
1
0
12.RunEn, 0
0.En, 8
TTCCLK
EXTCLK
SACLK
1.Cmd, 12
1
2
Fail-over Clock MUX
ENSACLK
ENINTCLK
FPGA2
Clock Detect
Clock Detect
Clock Detect
6. Stat, 9
6. Stat, 8
TTCCLK-IN
EXTCLK-IN
0
1
TTCREADY-IN
TTCCLK-OK
EXTCLK-OK
1
2
Fail-over Clock MUX
QPLL_LOCKED
Notes: TIM-3 CLOCK SWITCHING :
To ensure a safe and un-interruptible supply of 40 MHz clocks to all SCT FE modules, there a fail-safe network of two Zero Delay Glitch-Free Clock multiplexers/PLLs ( ICS 581—02 ).
The first multiplexer select between the X-tal generated internal clock or the externally supplied clock, with the X-tal clock being the default.
The second and final multiplexer is used to select the TTC-generated or stand-alone clock. If there is a failure of the TTC clock, this device switches automatically to the stand-alone clock within 3 clock periods.
Failure of any selected clock is also indicated by the de-assertion
of the TIM-OK line on the backplane.
TIM - OK :
"TIM-OK" is a +5V TTL “High-when-OK" output, and is set low by either the failure of the TTC clock as received by TIM, or also by the failure of the TIM X-tal clock oscillator when in stand-alone mode.
Additionally, "TIM-OK" will also be set low in case of failure of the +3V3 and/or +5V0 supply on TIM ( eg. by blown fuses ).
( "INTBUSY" & "OA_BUSY" in the Status Register 6 bit 2 do NOT get asserted when "TIM-OK" goes LOW = ie. is not OK )
( "ROD_CRATE_BUSYOUT" is NOT asserted when "TIM-OK" goes LOW = ie. is not OK )
DEFAULT :
After a switch-on and after each "O/A Reset", the internal X-tal clock is automatically selected to produce "CLKIN"
"TIM-OK" output is set to '1', and "TIM-OK" bit 13 in the Status Register 6 is also set to "1"
ERROR CONDITIONS :
1) when the "SetRunMode" bit 12 in Commands Register 1 is set to '1' :
if "TTCClkOn" line, AND/OR the "TTC-READY_IN" line, AND/OR the “QPLL_LOCKED” line go to '0', ( ie. The TTC clock is not present for more than 4 clock periods, or the TTCrq is not in a correct state ) :
Set "TIM-OK" output to '0', and set "TIM-OK" bit 13 in the Status Register 6 to "0"
When this 'error' conditions end, the "TIM-OK "output and "TIM-OK" bit 13 in the Status Register 6, should be reset back to "1"
2) when the "SetRunMode" bit 12 in Commands Register 1 is set to '0', AND when "EnExtClk" bit 8 in Commands Register 1 is set to '1' :
- if "SACLKON" line goes to '0' ( ie. the "CLKIN" is not present, due to either the absence of any external clock on inputs, or any of the inputs being set high, for more than 4 clock periods ) :
Set "TIM-OK" output to '0', and set "TIM-OK" bit 13 in the Status Register 6 also to "0"
When this 'error' conditions end, the "TIM-OK "output and "TIM-OK" bit 13 in the Status Register 6, should be reset back to "1”

## Slide 14: TIM Busy Signals Flow
MRMW/MP v1.2 17-06-04
Front Panel Busy Input
(NIM)
ECLEXTBUSY
EXTRODBUSYIN
EXT_BUSY
EN_EXT_BUSY
INT_BUSY
EN_INT_BUSY
EN_RODBUSY
RODBUSY
V_BUSY
BUSYOUT
LK6
PL41
MEXTBUSYOUT
RODBUSY(15)
RBmask
RB_MASK(15)
RODBUSY(14)
RB_MASK(14)
RODBUSY(0)
RB_MASK(0)
RODBUSY
ROD_CRATE_BUSYOUT
V_RODBUSY
Cmd, 8
SAMODE
Run En, 8
EN_EXTRODBUSY
LK4
LK1
PL167
NIMEXTBUSY
Stand-Alone System
Test
Busy
Signal
Deadtime
Status, 2
Status, 14
Status, 3
RBmask
RBmask
Enables, 15
Enables, 7
Cmd, 7
TIMOUTEN (enable for slot 13 only)
RUNMODE
Cmd,12
Run En, 7
Front Panel Busy Input
(ECL)
Backplane (J3) ROD Busy Inputs
LK7
1
4
Status, 0
Front Panel TIM Busy
LED (‘TB’)
Status, 1
Front Panel TIM Busy
(NIM)
Front Panel TIM Busy
(ECL)
Front Panel ROD Busy
LED (‘RB’)
Front Panel ROD Busy
Output (NIM/TTL)
Backplane ROD Busy
Output (NIM/TTL)
PLD/FPGA
Status, 7
Burst
Busy
Status, 4
1
1
1
1
Notes: DESCRIPTION OF BUSY's on TIM-3 :
1) On FPGA2, there are internal busys, which get asserted while any of the internally-generated TTC(x) functions are being generated.
They should be as long as the SCT serial code for each 'fast command' :
L1A : 3 clocks
BCR, FER, ECR, Spare : 7 clocks
CAL : 27 + pipeline-delay-to-trigger + 3 clocks
2) "OA_BUSY" is an OR of all above BUSYs and of "TESTBUSYON"
( "OA_BUSY" in the Status Register 6 bit 2 should NOT get asserted when "TIM-OK" goes LOW = ie. not OK )
3) "TESTBUSYON" is a busy line which is enabled by Commands Register 1 bit 13 "EnTestBusy", and cleared by Commands Register 1 bit 14 "ClrTestBusy" and/or by "OA_RESET"
"TESTBUSYON" is set by the next L1A after being enabled, and should prevent any internally-generated or externally-received fast commands after this L1A being acted upon
"TESTBUSYON" is also output to the Status Register 6 bit 14
4) "INTBUSY" is the "OA_BUSY" shown in the Status Register 6 bit 2
( "INTBUSY" & "OA_BUSY" in the Status Register 6 bit 2 should NOT get asserted when "TIM-OK" goes LOW = ie. not OK )
5) "VBUSY" is set in Commands Register 1 bit 7
6) "EXTBUSY" is the OR of externally-input "NIMEXTBUSY" and "ECLEXTBUSY" from the front panel
It is also a bit 0 in Status Register 6
7) Front panel signal "NIMEXTBUSYIN" is enabled onto the "NIMEXTBUSY" line by selecting link LK7 pins 1+2
( NOTE : If this selection in is used, the "EXTRODBUSYIN" should be grounded by selecting link LK7 pins 3+4 )
8) Front panel signal "NIMEXTBUSYIN" is enabled onto the "EXTRODBUSYIN"
line by selecting link LK7 pins 2+3
9) "BUSYOUT" is the OR of
a) "OA_BUSY" ( same as "INTBUSY" ) enabled by "EnIntBUSY" in Enables Register 0 bit 7
b) "EXTBUSY" enabled by "EnExtBUSY" in Enables Register 0 bit 15
c) "VBUSY" from Commands Register 1 bit 7
d) "RODBUSY", enabled by "EnRODBUSY" in TTC Enable Register 12 bit 7
It is also a bit 3 in Status Register 6
10) "BUSYOUTB" is the buffered "BUSYOUT" sent to 'TIM BUSY' LED
11) "NIMBUSYOUT" is a front panel output of the "BUSYOUTB" signal ( buffered version of "BUSYOUT" ), and can be selected by link LK6 to be either +ve or -ve going
12) "ECLBUSYOUT" is a front panel output, which is selected by link PL41 to be either "BUSYOUTB" signal as above ( PL41 pins 1+2 ), or the "MEXTBUSYOUT" signal ( PL41 pins 2+3 )
13) "MEXTBUSYOUT" is the "EXTBUSY" signal enabled by "EnExtBUSY" in Enables Register 0 bit 15
It is also a bit 1 in Status Register 6
14) "VRODBUSY" is set in Commands Register 1 bit 8
15) "RODBUSYOUT" is a front panel output, which is selected by link LK1 to be either NIM-type signal "ROD_CRATE_BUSYOUTB3" ( LK1 pins 2+3 ) – which can be selected either +ve or -ve going by link LK4 - or TTL-type signal ( LK1 pins 1+2 )
( "ROD_CRATE_BUSYOUT" should NOT get asserted when "TIM-OK" goes LOW = ie. is not OK )
( it should BE asserted while the TIM is in "SAMODE" )
16) "ROD_CRATE_BUSYOUT" is the OR of
a) masked OR of all the ROD-BUSY(x) signals received from TIM backplane
( see ROD Busy Register 16 and ROD Mask Register 15 )
b) "EXTRODBUSYIN" input from the front panel ( see par.8 ) and enabled by the "EnRODBUSY" bit 7 the TTC Enables Register 12
c) "VRODBUSY from Commands Register 1 bit 8
d) "SAMODE" from Status register 6 bit 11
17) "VRODBUSY" is set in Commands Register 1 bit 8
18) "RODBUSYOUT" is a front panel output, which is selected by link LK1 to be either NIM-type signal "ROD_CRATE_BUSYOUTB3" ( LK1 pins 2+3 ) – which can be selected either +ve or -ve going by link LK4 - or TTL-type signal ( LK1 pins 1+2 )
( "ROD_CRATE_BUSYOUT" should NOT get asserted when "TIM-OK" goes LOW = ie. is not OK )
( it should BE asserted while the TIM is in "SAMODE" )
19) "ROD_CRATE_BUSYOUT" is the OR of
a) masked OR of all the ROD-BUSY(x) signals received from TIM backplane
( see ROD Busy Register 16 and ROD Mask Register 15 )
b) "EXTRODBUSYIN" input from the front panel ( see par.8 ) and enabled by the "EnRODBUSY" bit 7 the TTC Enables Register 12
c) "VRODBUSY from Commands Register 1 bit 8
d) "SAMODE" from Status register 6 bit 11
20) "RODBUSYOUT" is the "ROD_CRATE_BUSYOUT" set in bit 7 of Status Register 6
21) "ROD_CRATE_BUSYOUTB1" is the "ROD_CRATE_BUSYOUT" sent to TIM backplane J3 pin 21D
22) "ROD_CRATE_BUSYOUTB2" is sent to the 'ROD BUSY' LED
23) BurstBUSY, SeqBUSY, SinkBUSY - see Status Register 6