Current  Status  of  ZEUS-MVD  “C+C  MASTER”  and  “C+_C  SLAVE”  modules

DESY,   15  February  2000,  updated to 14 March 2000

( MP-UCL )

1)	C+C  MASTER :				- Both modules fully tested stand-alone and with

the GFLT Simulator at UCL

2)	C+C  SLAVE :				- Two modules fully tested stand-alone and as a

system with C+C MASTER and the GFLT        Simulator at UCL

3)	C+C MASTER with two C+C SLAVEs :	-  System fully tested in one VME crate – both

stand alone and with the GFLT Simulator at           UCL

-  Soak testing done with the GFLT Simulator

using PC-driven software test programme to  check the  correct handling of Triggers, FLTNs, GBCNs and Busys shows no errors after runs lasting many millions of ‘events’

4)	TO BE DONE :				- Testing with  real GFLT at DESY

( this requires new cables to be manufactured )

5)	GFLT -> VETO WALL CABLE :		-  Testing with ~ 75 m of new, shielded, zero-

halogen “Twist’n’Flat” ribbon cable between   the GFLT simulator and the C+C MASTER shows that there is no actual loss of the phase lock between the GFLT  96nsec clock and the PLL clock regenerated on the C+C  MASTER .

This cable should be good enough for the ~52 m

-  Required to run between the GFLT in the

rucksack and the veto wall area electronics     racks

One C+C MASTER and one C+C SLAVE are now in DESY ready to be used in tests