## Slide 1: ABC130 reception tests
Bruce Gallop et al

## Slide 2: Test setup
HSIO (the tried and tested DAQ board)
  Generates/receives
    LVDS fast signals
    CMOS levels
Driver board
  FPGA with persistent program
  Can change signals as well
  Small (fast to program)
  Translate levels from LVDS to SLVS used by chip
Single chip card
  Passive components + ABC130

## Slide 3: Diced chips

## Slide 4: Dicing detail
Dicing not quite as
expected
Optimised for
TDC chip yield
To detriment of ABC130
Was another ABC130

## Slide 5: One chip on a test board

## Slide 6: Driver board

## Slide 7: Driver + single chip board at RAL

## Slide 8: Successes
Turn power on
  See currents (changes with clocks etc)
  Regulated voltage
Send commands to chip
  Correct de-multiplexing from L0
  Set registers
    The preamp current – Changes current draw (30-38mA)
    Direction bit – Changes from 0 to X
    Drive currents – Changes voltage swing
    Regulator voltage control
  Responds only when addressed by correct ID

## Slide 9: Powering
Offset from gnd for SLVS signals
Regulator output voltages
VDDD
GND
VDDA
GND

## Slide 10: Bidirectional signals
Left, with dir = 0
“0”
Not driven
XOFF
DATA

## Slide 11: Send command
Multiplex L0 and COM (40 into 80)
  Read register 0x20
1 0 1 1 1 1 1 0 1 1 1 1 1 1 1 1 1 1 0 1 0 0 0 0 0 0
Header Type HCC ABC Reg Read

## Slide 12: Unfortunately…
We don’t see data on the data line
Seems to be driving a constant 0

## Slide 13: What we’ve tried
Turn on clocks
Turn on regulators
Send command to turn on output FIFOs
  Without which no data is expected
Send command which generates data
  Read register (two varieties)
  R3
  L1
  Data from neighbour
  Each have different FIFOs

## Slide 14: Multiplexer
Write Reg 0x7 = 0x4010
(Preamp)
  0 = 26mV
  4 = 31
  8 = 36
  12 = 41
  16 = 46
  20 = 51
  24 = 56
  28 = 61
  31 = 65
<- With scope
Redo with DVM
0 = 56
31 = 96

## Slide 15: Other channels
Channel 6: 678mV VCS (nominal 900)
Channel 7: 751mV VCD (nominal 600)
Channel 8: 756mV VCSP (nominal 600)
Channel 9: 309mV VBASE (nominal 500)
Channel 10: 279mV VB (nominal 450)
Channel 11: 37mV VBG (nominal 592)
    Oops, testres is wrong

## Slide 16: Fix band gap voltage
Disconnect testres from driver board
  6: 678
  7: 748mV
  8: 753
  9: 309
  10: 280
  11: 620mV

## Slide 17: Scan Test data
Set scan enable
  Bit pattern comes out of SDO
  Longer pattern on BC chain (as expected)
Data clock scan chain (at least the part with 1s in it) is short enough that it’s easy to see changes directly on the scope
  Can also capture on on HSIO
Things that have an effect
  Read out (L1, R3, reg), all change same bit
  Data direction
  XOFF